REAGEN™

Flumequine ELISA Test Kit

RND99042

### Product Description

REAGEN™ Flumequine ELISA Test Kit enables government agencies, food manufacturers and quality assurance organizations to detect flumequine in various sample types and to satisfy customer concerns about food safety. REAGEN™ Flumequine  ELISA Test Kit provides a unique extraction method, which is accurate, rapid, cost effective and easy to use.

The unique features of the kit are:

- The extraction of Flumequine from feed, meat, tissue, serum and urine can be finished with a high recovery rate of 75 - 95%.

- A quick ELISA assay (less than 2 hours regardless of number of samples).

- High reproducibility.

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The flumequine antibody has been coated in the plate wells. During the analysis, sample is added along with the flumequine-horseradish peroxidase (Flumequine-HRP) conjugate. If the flumequine residue is present in the sample, it will compete for the flumequine antibody, thereby preventing the Flumequine-HRP from binding to the antibody attached to the well. The resulting color intensity,after addition of the HRP substrate (TMB), has an inverse relationship with the flumequine residue concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Flumequine ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8°C. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Flumequine Ab-coated Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8°C |
| Flumequin Standards: Negative control (white cap tube) 1 ng/mL (yellow cap tube) 5 ng/mL (orange cap tube) 10 ng/mL (pink cap tube) 25 ng/mL (purple cap tube) 75 ng/mL (blue cap tube) 1000 ng/mL (spiking, optional, red cap tube) | 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL | 2-8°C |
| Flumequine-HRP Conjugate | 8 mL | 2-8°C |
| 10X Sample Extraction Buffer ** | 25 mL |  |
| 20X Wash Solution ** | 30 mL |  |
| Stop Buffer ** | 14 mL |  |
| TMB Substrate ** | 12 mL |  |

*If you are not planning to use the kit for over 3 months, store HRP-Conjugate at -20°C preferably or in a freezer.

Sensitivity (Detection Limit)

| Sample Type | Detection Limit (ng/g or ppb) (Organic Reagent Extraction Method) | Detection Limit (ng/g or ppb) (Rapid Method) |
| --- | --- | --- |
| Feed | 1 | 10 |
| Meat/Liver/Kidney/Fish/Shrimp | 0.5 | 10 |
| Milk | 1 | 5 |
| Serum/Urine | 1 | 5 |

| Specificity (Cross-Reactivity) Fluoroquinolines | Cross-Reactivity (%) |
| --- | --- |
| Flumequine | 100 |
| Oxolinic acid | 25 |
| Pipemidic acid | 16 |
| Ofloxacin | 35 |
| Ciprofloxacin | 125 |
| Enrofloxacin | 112 |
| Cinoxacin | 2 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Ethyl Acetate

- Anhydrous Sodium Sulfate

### Warnings and Precautions

- The standards contain flumequine. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. HRP CONJUGATES AND PLATES ARE KIT-AND LOT-SPECIFIC.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting reagents into an empty microtiter plate, place the pipette tips

in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

1.   Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water

2.   Preparation of 35% Methanol/Sample Extraction Buffer

Mix 6.5 volumes of 1X Sample Extraction Buffer with 3.5 volumes of 100% methanol

Feed

Rapid Method

- To a reasonable amount of sample, add 5 times the amount of 70% methanol (e.g. 1 g of sample + 5 mL of 70% methanol); homogenize the sample with a suitable mixer.

- Take out 1.5 mL of the homogenized sample and centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

- Transfer 0.5 mL of the supernatant to a tube, add 0.5 mL of 1X Sample Extraction Buffer, and mix well.

- Use 100 µL of the sample per well for the assay.

Note: Dilution factor: 10.

Organic Reagent Extraction Method

- Homogenize the feed sample with a suitable mixer.

- Weigh out 2.0 g of the feed sample, add 10 mL of ethyl acetate to the sample, shake 20 min with a mechanical shaker.

- Centrifuge for 5 minutes at 4,000 x g at room temperature  (20 – 25°C / 68 – 77°F).

- Transfer 5 mL of the ethyl acetate supernatant (corresponding to 1 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 50-60°C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50-60°C water bath.

- Dissolve the dried residue in 2 mL of n-hexane.

- Add 1 mL of 35% Methanol/Sample Extraction Buffer and mix by vortexing at maximum speed for 2 minutes.

- Centrifuge for 10 minutes at 4,000 x g at room temp. Discard the upper hexane layer.

- Use 100 µL of the lower aqueous layer per well for the assay.

Note: Dilution factor: 1.

Meat/Liver/Kidney/Fish/Shrimp

Rapid Method

- Remove fat from the sample. Homogenize the sample with a suitable mixer.

- Weigh out 1.0 g of the homogenized sample and add 4 mL of 70% methanol.

- Vortex for 10 minutes at maximum speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

- Transfer 0.5 mL of the supernatant to a tube, add 0.5 mL of 1X Sample Extraction Buffer, and mix well.

- Use 100 µL for the assay.

Note: Dilution factor: 10.

Organic Reagent Extraction Method

- Remove fat from the sample. Homogenize a reasonable amount of sample with a suitable mixer.

- Weigh out 3 g of the homogenized sample and mix with 6 mL of ethyl acetate, vortex 3 min at max speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

- Transfer 4 mL of the ethyl acetate supernatant (corresponding to 2 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 60-70°C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 60-70°C water bath.

- Dissolve the dried residue in 2 mL of n-hexane.

- Add 1 mL of 35% Methanol/Sample Extraction Buffer and mix by vortexing at maximum speed for 2 minutes.

- Centrifuge for 10 minutes at 4,000 x g at room temp. Discard the upper hexane layer.

- Use 100 µL of the lower aqueous layer per well for the assay. In case emulsion happened, remove the upper n-hexane layer and incubate the lower aqueous layer in water bath for 3 min at 80-95°C.

Note: Dilution factor: 0.5.

Milk

Rapid Method

- For fat-free milk, take out 200 µL of the milk sample, add 800 µL of 35% Methanol/Sample Extraction Buffer, and mix well. Take 100 µL of the diluted sample per well for the assay.

- For regular milk with fat, centrifuge 1 mL of the milk sample at 4,000 x g for 5 minutes, discard the upper fat layer. Take out 200 µL  of  the milk  sample, add 800 µL  of  35% Methanol/Sample Extraction Buffer, and mix well. Take 100 µL of the diluted sample per well for the assay.

Note: Dilution factor: 5.

Organic Reagent Extraction Method

- Weigh out 2.0 g of milk sample and add 4 g of anhydrous sodium sulfate. Homogenize the sample with a suitable mixer.

- Add 10 mL of ethyl acetate to the sample, shake 20 min with a mechanical shaker.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

- Transfer 5 mL of the ethyl acetate supernatant (corresponding to 1 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 50-60°C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50-60°C water bath.

- Add 1 mL of 35% Methanol/Sample Extraction Buffer and mix by vortexing at maximum speed for 2 minutes.

- Use 100 µL of the sample per well for the assay.

Note: Dilution factor: 1.

Serum/Urine

Rapid Method

- Centrifuge 0.5 mL of urine at 4,000 x g for 5 minutes.

- Take out 0.2 mL of the supernatant, add 0.8 mL of 35% Methanol/Sample  Extraction Buffer.

- Use 100 µL per well for the assay.

Note: Dilution factor: 5.

Organic Reagent Extraction Method

- Weigh out 2.0 g of sample and add 4 g of anhydrous sodium sulfate. Homogenize the sample with a suitable mixer.

- Add 10 mL of ethyl acetate to the sample, shake 20 min with a mechanical shaker.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

- Transfer 5 mL of the ethyl acetate supernatant (corresponding to 1 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 50-60°C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50-60°C water bath.

- Add 1 mL of 35% Methanol/Sample Extraction Buffer and mix by vortexing at maximum speed for 2 minutes.

- Use 100 µL of the sample per well for the assay.

Note: Dilution factor: 1.

FLUMEQUINE ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

1.  Preparation of  1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

| Label the individual strips that will be used and aliquot reagents as the following example: Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| 1X HRP-Conjugated | 50 µL | 1.2 mL |
| 1X Wash Solution | 1.0 mL | 24 mL |
| Stop Buffer | 100 µL | 2.4 mL |
| TMBSubstrate | 100 µL | 2.4 mL |

1. Add 100 µL of each Flumequine Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

2. Add 100 µL of each sample in duplicate into different sample wells.

3. Add 50 µL of 1X HRP-Conjugated . and mix well by gently rocking the plate manually for 1 minute

4. Incubate the plate for 60 minutes at room temperature (20 – 25°C / 68 – 77°F) (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

5. Wash the plate 3 times with 250 µL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

6. Add 100 µL of TMB substrate. Time the reaction immediately after adding the substrate.Mix the solution by gently rocking the plate manually for 1 minute while incubating

7. After incubating for 20 minutes at room temperature (20 – 25°C / 68 – 77°F), add 100 µLof Stop Buffer to stop the enzyme reaction.

8. Read the plate as soon as possible following the addition of Stop  Buffer on a plate reader with 450 nm wavelength (Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Flumequine Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical Flumequine standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us