REAGEN™

Trimethoprim ELISA Test Kit

RND99037

### Product Description

REAGEN™Trimethoprim ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Trimethoprim (TMP) in fish, shrimp, feed, meat (beef, chicken and pork), milk,serum and urine.

The unique features of the kit are:

- Rapid extraction method for various sample types

- High recovery rates (>80%)

- High sensitivity (0.015 ng/g or ppb)

- Quick ELISA assay (less than 2 hours regardless of number of samples)

- High reproducibility

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The plate wells are coated with Trimethoprim (TMP) antibody. During analysis, sample is added along with the TMP-horseradish peroxidase (TMP-HRP) conjugate. If the TMP residue is present in the sample, it will compete for the TMP antibody, thereby preventing the TMP-HRP from binding to the antibody attached to the well. The resulting color intensity, after addition of the HRP substrate (TMB), has an inverse relationship with the TMP residue concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Trimethoprim ELISA Test Kit has the capacity for 96 determinations of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Trimethoprim Plate | 1x96-well Plate(8wellsx12 strips) | 2-8C |
| Trimethoprim Standards: Negative control (white cap tube) 0.015 ng/mL (yellow cap tube) 0.04 ng/mL (orange cap tube) 0.12 ng/mL (pink cap tube) 0.36 ng/mL (purple cap tube) 1.08 ng/mL (blue cap tube) 1000 ng/mL(spiking, optional, red cap tube) | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | 2-8C 2-8C |
| TMP-HRP Conjugate | 13mL | 2-8C |
| 10X Sample Extraction Buffer | 25 mL |  |
| 20X Wash Solution ** | 28 mL |  |
| Stop Buffer ** | 14 mL |  |
| TMB Substrate ** | 12 mL |  |
| Milk Extraction Buffer(Optional) | 25mL |  |

* If you are not planning to use the kit for over 1 month, store Trimethoprim standards and Trimethoprim -HRP Conjugate at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Feed | 0.8 |
| Fish/Shrimp/Meat/Liver/Kidney | 0.2 (Method I); 10 (II) |
| Urine /Serum/ Plasma | 0.2 |
| Milk/milk powder | 10 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Trimethoprim | 100 |
| Sulfapyridine | <0.1 |
| Sulfamethoxypyridazine | <0.1 |
| Sulfadiazine | <0.1 |
| Sulfamoxole | <0.1 |
| Sulfisoxazole | <0.1 |
| Sulfathiazole | <0.1 |
| Sulfamerazine | <0.1 |
| Sulfaquanidine | <0.1 |
| Sulfadoxine | <0.1 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni Tissue Master Homogenizer)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- N-hexane

- Methanol

### Warnings and Precautions

- The standards contain trimethoprim. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. HRP CONJUGATE AND PLATES ARE KIT-AND LOT-SPECIFIC.

- Try to maintain a laboratory temperature of 20°–25°C. Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

1.Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

2. The recommended dilution factors for different type of samples shown below are minimum ones in order to achieve the lowest detection limit. The dilution factor can be tailored to meet larger sample positive cutoff value.

### Feed

    - Homogenize a reasonable amount of feed sample using a suitable mixer.

    - Add 20 mL of 0.1 M HCl to 2 g of the homogenized sample. Vortex for 15 minutes on a multi-tube vortexer or shaker.

    - Centrifuge for 10 minutes at 2,000 x g at room temperature (20 – 25°C), transfer 1 mL of the supernatant to a new tube, and add ~55 µL of 1 M NaOH to adjust the pH to 6-8 (the amount of 1 M NaOH can be adjusted for different feed samples).

    - Centrifuge the sample for 10 minutes at 2,000 x g, transfer 0.5 mL of the supernatant to a new tube, add 0.5 mL of 1X Sample Extraction Buffer, mix well.

    - Use 50 µL of the sample for the assay.

Note: Dilution factor: 20

### Fish/Shrimp/Meat/Liver/Kidney

Method I(For low concentration detection)

    - Weigh 2 g homogenized meat sample and add 6 mL of methanol and 2 mL of hexane. Vortex for 3 minutes at maximum speed or shake for 20 minutes on a multi-tube vortexer or shaker.

    - Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F), remove the hexane layer completely and then transfer 0.4 mL of the supernatant to a new tube.

    - Use a rotary evaporator to dry the sample in a 50-60°C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50-60°C water bath.

    - Add 500µL of 1X Sample Extraction Buffer and 500µL of hexane, vortex vigorously for 1 minute at maximum speed.

    - Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C), remove the hexane layer completely, use 50µL of the lower layer sample for the assay.

Note: Dilution factor:5. If needed, the sample obtained from step 5 can be diluted with 1X Sample Extraction Buffer.

Method II (For high concentration detection)

    - Weigh 1g homogenized meat sample and add 6 mL of methanol and 2 mL of hexane. Vortex for 3 minutes at maximum speed or shake for 10minutes on a multi-tube vortexer or shaker.

    - Centrifuge for 10 minutes at 4,000 x gat room temperature (20 – 25°C), remove the hexane layer completely and then transfer 20 µL of the supernatant to a new tube.

    - Add 980µL of 1X Sample Extraction Buffer,vortex vigorously for 30seconds.

    - Use 50 µL of the sample for the assay.

Note:Dilution factor: 350.

### Milk

    - Add 1 mL of milk sample into a plastic tube. Add 2 mL of distilled water and 300 µL of Milk Extraction Buffer to each sample.

    - Vortex vigorously for 60 seconds.

    - Centrifuge at 4,000 x g for 10 minutes.

    - Take 20 µL of supernatant and add 1980 µL of 1X Sample Extraction Buffer, mix well.

    - Use 50 µL of the sample for the assay.

Note: Dilution factor: 330.

### Milk Powder

    - Weigh out 1 g of milk powder in a centrifugal glass vial, add 9 mL of distilled water and dissolve by shaking.

    - Take out 1 mL of the dissolved milk sample, add 100 µL of Milk Extraction Buffer to each sample. Vortex vigorously for 60 seconds.

    - Centrifuge at 4,000 x g for 10 minutes.

    - Take 30 µL supernatant and add 870 µL of 1X Sample Extraction Buffer, mix well.

    - Use 50 µL of the sample for the assay.

Note: Dilution factor: 330.

### Urine/ Serum/ Plasma

- Centrifuge 0.5 mL of the urine sample at 4,000 x g for 5 minutes.

- Take out 50 L of the supernatant, add 200 L of 1X Sample Extraction Buffer, and mix well.

- Use 50 L per well for the assay.

Note: Dilution factor: 5. If needed, the sample can be further diluted with 1X Sample Extraction Buffer.

TRIMETHOPRIM ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| TMP -HRP Conjugate | 100 L | 2.4 mL |
| 1X Wash Solution | 1.0 mL | 24 mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

- Add 50 L of each trimethoprim Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration). Add 50 L of each sample in duplicate into different sample wells.

- Add 100 L of TMP -HRP Conjugate to each well and mix well by gently rocking the plate manually for 1 minute

- Incubate the plate for 1 hour at room temperature (20 – 25C / 68 – 77F).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Don’t allow the plate to air dry between working steps).

- Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating ( Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Trimethoprim Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical TMP standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com   reagenllc@gmail.com

Web.: www.reagen.us