REAGEN™

Difloxacin ELISA Test Kit

RND99078

### Product Description

REAGEN™Difloxacin ELISA Test Kit enables government agencies, food manufacturers and quality assurance organizations to detect major Difloxacin in various sample types and to satisfy customer concerns about food safety. Difloxacin ELISA Test Kit iprovides a unique extraction method, which is accurate, rapid, cost effective and easy to use. .

The unique features of the kit are:

- Extraction of difloxacins from chicken, fish/shrimp and milk with a high recovery rate.

- High reproducibility.

- A quick ELISA assay (50 minutes incubation time).

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The difloxacin antibody  has been coated in the plate wells. During the analysis, sample is added along with the difloxacin-horseradish peroxidase (Difloxacin-HRP) conjugate. If the difloxacin residue is present in the sample, it will compete for the difloxacin antibody, thereby preventing the Difloxacin-HRP from binding to the antibody attached to the well. The resulting color intensity, after addition of the HRP substrate (TMB), has an inverse relationship with the difloxacin residue concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Difloxacin ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with desiccant provided in the original package. Store the kit at 2-8°C*. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Difloxacin-coated Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8°C |
| Difloxacin Standards: Negative control (white cap tube) 0.05 ng/mL (yellow cap tube) 0.2 ng/mL (orange cap tube) 0.8 ng/mL (pink cap tube) 3.2 ng/mL (purple cap tube) 12.8 ng/mL (blue cap tube) 500 ng/mL (spiking, optional, red cap tube) | 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL | 2-8°C |
| Difloxacin-HRP | 12 mL | 2-8°C |
| 20X Wash Solution ** | 28mL |  |
| Stop Buffer ** | 14 mL |  |
| TMB Substrate ** | 12 mL |  |
| 10X Sample Extraction Buffer** | 25 mL |  |
| Meat Extraction Buffer I (Optional) | 5mL |  |
| Meat Extraction Buffer II (Optional) | 5 mL |  |

*If you are not planning to use the kit for over 1 months, store  Difloxacin-HRP at -20°C or in a freezer.

Sensitivity (Detection Limit)

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Egg | 1.0 |
| Honey | 1.0 |
| Feed | 1.0 |
| Fish/Shrimp | 1.0 |
| Meat/Liver/Kidney | 1.0 |
| Milk | 0.5 |
| Serum | 0.5 |
| Urine | 0.5 |

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Difloxacin | 100 |
| Enoxacin | 18 |
| Pipemidic acid | 15 |
| Benofloxacin | 7 |
| Oxolin acid | 6 |
| Nalidixic acid | 4 |

Specificity (Cross-Reactivity)

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

### Warnings and Precautions

- The standards contain difloxacin. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. HRP CONJUGATES AND PLATES  ARE KIT-AND LOT-SPECIFIC.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

1.	Preparation of 1X Sample Extraction Buffer

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

2.	Preparation of 35% Methanol/Sample Extraction Buffer

Mix 6.5 volumes of 1X Sample Extraction Buffer with 3.5 volumes of 100% methanol.

Notes: The recommended dilution factors for different type of samples shown below are minimum ones in order to achieve the lowest detection limit. The dilution factor can be tailored to meet larger sample positive cutoff value. For example, if the positive cutoff value for meat sample is set up at 100 ppb, dilute meat sample with 1X Sample Extraction Buffer 100 fold instead of 10. After the dilution, the positive sample cutoff value stays in the linear range of the standard curve.

Egg

    - Weigh out 1.0 g of egg sample and add 4 mL of 70% methanol.

    - Vortex for 2 minutes at maximum speed.

    - Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

    - Transfer 0.5 mL of the supernatant to a tube, add 0.5 mL of 1X Sample Extraction Buffer, and mix well.

    - Use 50 µL of the sample per well for the assay.

Note: Dilution factor: 10.

Feed

1.  To a reasonable amount of sample, add 5 times the amount of 70%  methanol (e.g. 1 g of sample + 5 mL of 70% methanol); homogenize the sample with a suitable mixer.

2.  Take out 1.5 mL of the homogenized sample and centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

3.  Transfer 0.5 mL of the supernatant to a tube, add 0.5 mL of 1X Sample Extraction Buffer, and mix well.

4.  Use 50 µL of the sample per well for the assay.

Note: Dilution factor: 10.

Fish/Shrimp

- Remove fat from the sample. Homogenize the sample with a suitable mixer.

- Weigh out 1.0 g of the homogenized sample and add 4 mL of 70% methanol.

- Vortex for 10 minutes at maximum speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 0.5 mL of the supernatant to a tube, add 0.5 mL of 1X Sample Extraction Buffer, and mix well.

- Use 50 μL for the assay.

Note: Dilution factor: 10. If the expected detection range is 10-50 ppb, add one more dilution step after Step 5 : “ Take out 0.2 mL of the sample, add 0.6 mL of 35% methanol/ Sample Extract ion Buff er, mix wel l ”. The dil ution factor willbe 40.

Honey

- Take out 0.1 g of honey sample, add 1.9 mL of 35% Methanol/Sample Extraction Buffer and mix well.

- Use 50μL per well for the assay.

Note: Dilution factor: 20. If needed, the sample can be further diluted with 35% Methanol/Sample Extraction Buffer.

Meat/Liver/Kidney

- Remove fat from the sample. Homogenize the sample with a suitable mixer.

- Weigh out 1.0 g of the homogenized sample and add 4 mL of 35% methanol/Sample Extraction Buffer and 50μL of Meat Extraction Buffer I..

- Vortex for 10 minutes at maximum speed.

- Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

- Transfer 0.5 mL of the supernatant to a tube, add 25μL of Meat Extraction Buffer II and 1.5 mL of 35% Methanol/Sample Extraction Buffer, vortex for 30 seconds.

- Use 50 µL for the assay.

Note: Dilution factor: 20.

Milk

- For fat-free milk, take out 200 µL of the milk sample, add 800 µL of 35% Methanol/Sample Extraction Buffer, and  mix well. Take 50 µL of the diluted sample per well for the assay.

- For regular milk with fat, centrifuge 1 mL of the milk sample at 4,000 x g for 5 minutes, discard the upper fat layer. Take out 200 µL  of  the milk  sample, add 800 µL  of  35% Methanol/Sample Extraction Buffer, and mix well. Take 50 µL of the diluted sample per well for the assay.

Note: Dilution factor: 5.

Serum

- Centrifuge 0.5 mL of urine at 4,000 x g for 5 minutes.

- take out 0.2 mL of the supernatant, add 0.8 mL of 35% Methanol/Sample  Extraction Buffer.

- Use 50 µL per well for the assay.

Note: Dilution factor: 5. If needed, the sample can be further diluted with 35% Methanol/Sample Extraction Buffer.

Urine

- Centrifuge 0.5 mL of urine at 4,000 x g for 5 minutes.

- Take out 0.2 mL of the supernatant, add 0.8 mL of 35% Methanol/Sample Extraction Buffer.

- Use 50 µL per well for the assay.

Note: Dilution factor: 5

DIFLOXACIN  ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

1.Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Difloxacin-HRP | 100 µL | 2.4 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 µL | 2.4 mL |
| TMBSubstrate | 100 µL | 2.4 mL |

- Add 50 µL of each Difloxacin Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

- Add 50 µL of each sample in duplicate into different sample wells.

- Add 100 µL of Difloxacin-HRP and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 30 minutes at room temperature (20 – 25°C / 68 – 77°F).

- Wash the plate 3 times with 250 µL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 µL of TMB substrate. Time the reaction immediately after adding the substrate.Mix the solution by gently rocking the plate manually for 1 minute while incubating

- After incubating for 20 minutes at room temperature (20 – 25°C / 68 – 77°F), add 100 µLof Stop Buffer to stop the enzyme reaction. (Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the plate while incubating is recommended).

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength (Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Difloxacin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve. A special program with Excel functionality,

The following figure is a typical difloxacin standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us