REAGEN™

Doramectin ELISA Test Kit

RND99069

### Product Description

REAGEN™ Doramectin (Dectomax) ELISA Test Kit enables government agencies food manufacturers and quality assurance organizations to detect doramectin, abamectin and other avermectins in bovine muscle sample and to satisfy customer concerns about food safety. The REAGEN™ Doramectin ELISA Test Kit for Bovine Muscle provides a unique extraction method, which is accurate, rapid, cost effective and easy to use.

The unique features of the kit are:

- The extraction of Doramectin with a high recovery rate of 75 - 100%

- Rapid ELISA assay (less than 2 hours regardless of number of samples)

- High reproducility

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During analysis, sample is added along with the primary antibody specific for doramectin. If doramectin is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Doramectin (Dectomax) ELISA Test Kit for Bovine Muscle has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the  desiccant provided in the original package. Store the kit at 2-8°C*. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Doramectin-Coated Microtiter Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8℃ |
| Doramectin Standards: Negative control (white cap tube) 1.5 ng/mL (yellow cap tube) 5ng/mL (orange cap tube) 15 ng/mL (pink cap tube) 30 ng/mL (purple cap tube) 60 ng/mL (blue cap tube) 1200 ng/mL (spiking, optional, red cap tube) | 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL | 2-8℃* |
| Doramectin Antibody #1 | 12 mL | 2-8℃* |
| 100X HRP-Conjugated Antibody #2 | 250μL | 2-8℃* |
| Antibody #2 Diluent | 20 mL | 2-8℃ |
| 10×Sample Extraction Buffer | 25 mL |  |
| 20× Wash Solution | 28 mL |  |
| Stop Buffer | 14 mL |  |
| TMB Substrate | 12 mL |  |
| Doramectin Clean Up Mix | 25 g |  |

* If you are not planning to use the kit for over 1 month,store Doramectin Standaeds,Doramectin Antibody #1,and 100X HRP-Conjugated Antibody #2 in a freezer at -20°C.

### Sensitivity(Detection Limit)

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Bovine Muscle | 5 |

### Specificity(Cross-Reactivity)

| Molecules | Cross-Reactivity (%) |
| --- | --- |
| Doramectin | 100% |
| Emamectin | 100% |
| Ivermectin | 100% |
| Abamectin | 113% |
| Eprinomectin | 100% |
| Moxidectin | <10% |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer

- Vortex mixer

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Acetonitrile

### Warnings and Precautions

- The standards contain Doramectin. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not mix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT-AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes .

- Try to maintain a laboratory temperature of 20–25°C (68–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important to the assay.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration,as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2 - 4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25°C / 68 – 77°F) or in a refrigerator before use.

1.Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

2.Preparation of 35% Ethanol/ Sample Extraction Buffer

Mix 3.5 volumes of 100% ethanol and 6.5 volumes of 1X Sample Extraction Buffer

### Bovine Muscle(raw and cooked)

- Homogenize a reasonable amount of bovine muscle with a suitable mixer.

- Weigh out 2 g of the homogenized muscle sample into a 15 mL conical tube.

- Add 6 mL of acetonitrile to each tube containing meat.

- Shake vigorously for 60 seconds until muscle becomes fibrous and moves freely in solvent. Some samples might require additional shaking until fibrous. This increases the surface area of the beef and its contact with the solvent.

- Add 1 g of Anhydrous MgSO4 into one 15 mL tube and shake immediately for 5 seconds(Prevents conglomeration). Repeat this step for each sample tube.

- Once all tubes contain Anhydrous MgSO4, shake for an additional 60 seconds using a multi-tube vortex.

- Add 2 mL of Hexane to each sample tube and shake quickly to mix well (Approximately 5 seconds).

- Add all 15 mL conical tubes from step 4 into a multi-tube vortex and shake at maximum speed for 15 minutes.

- Centrifuge the sample for 10 minutes at 4,000xg at room temperature (20 – 25°C).

- Remove the Hexane layer completely and transfer 1.5 mL of acetonitrile to a 2 mL tube containing at least 375 mg of Doramectin Clean Up Mix.

- Vortex 1 minute in multi-tube vortex. Let samples rest 5 minutes.

- Centrifuge the sample for 5 minutes at 4,000xg at room temperature.

- Remove 1 mL of supernatant and add to a new 2 mL tube containing 50 µL of DMSO.

- Dry the samples using nitrogen gas with a TurboVap concentrator set at 7-10 psi and 50°C for 30 minutes duration or until only 50 µL remains.

- Add 450 μL of 35% Ethanol/Sample Extraction Buffer to the dried tube and vortex vigorously for 2 minutes.

- Use 50 µL for the ELISA.

Note: Dilution factor: 2.

** Alternative sample drying methods:

- The sample may be dried by blowing nitrogen gas in a 50C waterbath for 30 minutes or until only 50 µL remains.

- Take 1 mL from step 11 into a 2 mL plastic tube and dry in a miVac sample concentrator set at 50C for 30 minutes or until only 50 µL remains

DORAMECTIN ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

1.Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2 Diluent.

2.Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Doramectin Antibody #1 | 100 L | 2.4 mL |
| 1X-HRP Conjugate Antibody #2 | 150 L | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

- Add 50 L of each Doramectin Standard in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

- Add 50 L of each sample in duplicate into different sample wells.

- Add 100 L of Doramectin Antibody #1 to each well and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 150 L of 1X HRP-Conjugated Antibody #2. Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F) (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Doramectin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical Doramectin standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250     FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us