REAGEN™

Lincomycin ELISA Test Kit

RND99058

### Product Description

REAGEN™ lincomycin ELISA Test Kit enables government agencies, food manufacturers and quality assurance organizations to detect major lincomycin to as low as 0.5 ppb level in various sample types and to satisfy customer concerns about food safety. Lincomycin ELISA Test Kit provides a unique extraction method, which is accurate, rapid, cost effective and easy to use. .

The unique features of the kit are:

- The extraction of lincomycins from egg, honey, feed, meat, milk, tissue, serum and    urine can be finished in 10 - 30 minutes with a high recovery rate of 75 - 105%.

- High reproducibility.

- A quick ELISA assay (less than 2 hours regardless of number of  samples..

### Procedure Overview

The method  is  based on  a competitive  colorimetric ELISA  assay. The  lincomycin antibody  has been coated in  the plate wells.  During the analysis,  sample is  added along with  the lincomycin-horseradish peroxidase  (Lincomycin-HRP)  conjugate.  If the  lincomycin  residue  is  present  in  the  sample, it  will compete  for  the  lincomycin  antibody,  thereby  preventing  the  Lincomycin-HRP  from  binding  to the antibody attached  to the well.  The resulting color  intensity, after addition  of the HRP substrate  (TMB),has an inverse relationship with the lincomycin residue concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ lincomycin ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused micro wells to the foil bag and reseal them with desiccant provided in the original package. Store the kit at 2-8°C*. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Lincomycin Ab-coated Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8°C |
| Lincomycin Standards: Negative control (white cap tube) 0.1 ng/mL (yellow cap tube) 0.25 ng/mL (orange cap tube) 0.6 ng/mL (pink cap tube) 1.5 ng/mL (purple cap tube) 3.75 ng/mL (blue cap tube) 1000 ng/mL (spiking, optional, red cap tube) | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | 2-8°C |
| HRP-Conjugated | 6 mL |  |
| 200X Sample Extraction Buffer ** | 25 mL |  |
| 20X Wash Solution ** | 30 mL |  |
| Stop Buffer ** | 14 mL |  |
| TMB Substrate ** | 10 mL |  |
| Sample balance buffer | 0.5 ml |  |
| 10xPBS | 25 ml |  |

*If you are not planning to use the kit for over 3 months, store  lincomycin HRP-Conjugated  at -20°C preferably or in a freezer.

### Sensitivity (Detection Limit)

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Egg | 1.0 |
| Honey | 4.0 |
| Feed | 1.0 |
| Meat/Liver/Kidney/Fish/Shrimp | 1.0 |
| Milk | 2.0 |
| Milk powder | 5.0 |
| Serum | 1.0 |
| Urine | 1.0 |

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| lincomycin | 100 |
| Clindamycin | <0.1 |
| Erythromycin | <0.1 |
| Tylosin | <0.1 |
| Apramycin | <0.1 |
| Spectinomycin | <0.1 |
| TILM | <0.1 |

### Specificity (Cross-Reactivity)

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 uL pipettes

- Multi-channel pipette: 50-300 uL (Optional)

### Warnings and Precautions

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting reagents into an empty microtiter plate, place the pipette tips

in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

1.    Preparation of 1X Sample Extraction

Mix 1 volume of 200X Sample Extraction Buffer with 199 volumes of distilled water.

2.    Preparation of 1X PBS

Mix 1 volume of the 10X PBS with 9 volumes of distilled water.

3.    Preparation of 1X PBS-Tween Buffer ( for milk sample)

Mix 1 volume of the  10X PBS with 9 volumes of  distilled water, then add 0.1% Tween  20 (e.g. Add 0.1 g of Tween 20 in 100 mL of 1X PBS).

Egg

    - To a  reasonable amount  of the egg  sample, add  10 times  the amount  of 1X  Sample Extraction Buffer (e.g. 1 g of sample + 10  mL of 1X Sample Extraction Buffer), homogenize the sample with a suitable mixer.

    - Take out  1.2 mL  of the  homogenized sample and  centrifuge for  5 minutes  at 4,000  x g at room temperature (20 – 25°C / 68 – 77°F).

    - Transfer 0.5 mL of the supernatant to a new tube.

    - Incubate the supernatant at 75°C for 5 minutes and vortex for 1 min at maximum speed.

    - Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

    - Transfer 0.2  mL of  the supernatant to  a new  tube, add  5µL   of Sample  Balance Buffer and  mix well.

    - Use 100 µL per well for the assay

Note: Dilution factor: 10.

Feed

    - To a  reasonable amount  of the feed  sample, add  10 times  the amount  of 1X  Sample Extraction Buffer (e.g. 1 g of sample + 10  mL of 1X Sample Extraction Buffer), homogenize the sample with a suitable mixer.

    - Take out  1.2 mL  of the  homogenized sample and  centrifuge for  5 minutes  at 4,000  x g at room temperature (20 – 25°C / 68 – 77°F).

    - Transfer 0.5 mL of the supernatant to a new tube.

    - Incubate the supernatant at 75°C for 5 minutes and vortex for 1 min at maximum speed.

    - Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

    - Transfer 0.2  mL of  the supernatant to  a new  tube, add  5µL   of Sample  Balance Buffer and  mix well.

    - Use 100 µL per well for the assay

Note: Dilution factor: 10.

Honey

    - Weigh 0.5 g of honey into a centrifuge vial. Add 3.5 mL of 1X PBS.

    - Ensure the sample is fully dissolved in the PBS by heating for 5 – 10 minutes at 50 °C.

    - Vortex for 10 minutes in a multi-tube vortexer or shake 30 minutes on a shaker.

    - Centrifuge for 10 minutes at 4,000 x g.

    - Transfer 100µL  of the supernatant to a new tube containing 400µL  of 1X PBS.

    - Vortex for 1 minute.

    - Use100µL  per well in the assay.

Note:  Dilution factor:  40.  If needed,  the  sample  obtained from  step  6  can be  further diluted with 1X PBS.

Meat/Liver/Kidney/Fish/Shrimp

    - Remove fat from the sample. Homogenize the sample with a suitable mixer.

    - Add 4.5  mL  of 1X  Sample  Extraction Buffer  to 0.5  g  of the  homogenized  sample. Vortex  for  3 minutes at maximum speed or rotorack for 20 minutes ( e.g. VWR Tube Rotator  ).

    - Take out  1.2 mL  of the  homogenized sample and  centrifuge for  5 minutes  at 4,000  x g  at room temperature (20 – 25°C / 68 – 77°F).

    - Transfer 0.5 mL of the supernatant to a new tube (Avoid the fatty layer!).

    - Incubate the supernatant at 75°C for 5 minutes and vortex for 1 minute at max speed.

    - Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

    - Transfer 0.2 mL of the  supernatant to a new tube ( Avoid the fatty layer!),  add 5µL  of Sample Balance Buffer and mix well.

    - Use 100 µL for the assay.

Note: Dilution factor: 10.

Milk

- For fat-free milk, dilute the milk sample with 1X PBS-Tween Buffer (1:19)  (e. g. 20µL  of milk + 380µL of 1X PBS-Tween Buffer). Take100µL  of the diluted sample per well for the assay.

- For the  regular milk  with fat,  centrifuge  the milk  sample at  4,000 x  g for  5 minutes,  discard the upper fat layer. Dilute the sample with 1X PBS-Tween Buffer (1:19) (e. g. 20µL  of milk + 380µL  of 1X PBS-Tween Buffer). Take100µL  of the diluted sample per well for the assay.

Note: Dilution factor: 20. If needed, the sample can be further diluted with 1X PBS.

Milk Powder

- Weigh out  2 g  of the  milk powder  sample in  a centrifugal  vial, add  10 mL  of distilled  water and dissolve by shaking.

- Centrifuge at 4,000 x g for 5 minutes, discard the upper lipid layer.

- Dilute the  sample  with 1X  PBS-Tween Buffer  (1:9) (e.  g.  20 µL of the sample  + 180 µL of 1X PBS-Tween Buffer).

- Take 100µL  of the diluted sample per well for the assay.

Note: Dilution factor: 50. If needed, the sample can be further diluted with 1X PBS.

Serum  Urine

- Centrifuge 0.5 mL of the serum sample at 4,000 x g for 5 minutes.

- Take 20µL  of the supernatant and add 180µL  of 1X PBS, mix well.

- Use 100µL  per well for the assay.

Note: Dilution factor: 10. If needed, the sample can be further diluted with 1X PBS.

LINCOMYCIN ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

1.  Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2 Diluent.

2.  Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

3.	Preparation of 1X Sample Extraction Buffer

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

4.	Preparation of 35% Methanol/Sample  Extraction  Buffer (For tests on Milk, Serum and Urine samples only)

Mix 6.5 volumes of 1X Sample Extraction Buffer with 3.5 volumes of 100% methanol.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| HRP-Conjugated | 50 µL | 1.2 mL |
| 1X Wash Solution | 1.0 mL | 24 mL |
| Stop Buffer | 100 µL | 2.4 mL |
| TMB Substrate | 100 µL | 2.4 mL |

- Add 100 µL of each lincomycin Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

- Add 100 µL of each sample in duplicate into different sample wells.

- Add 50 µL of HRP-Conjugated. Incubate the plate for 30 minutes at room temperature (20 – 25°C / 68 – 77°F) (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 µL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 µL of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating

- After incubating for 15-20 minutes at room temperature (20 – 25°C / 68 – 77°F), add 100 µLof Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop  Buffer on a plate reader with 450 nm wavelength (Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### lincomycin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us