REAGEN™

Sulfaquinoxaline (SQX) ELISA Test Kit

RND99036

### Product Description

REAGEN™Sulfaquinoxaline ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Sulfaquinoxaline (SQX) in fish, shrimp, honey, kidney, liver, meat (beef, chicken and pork), milk, serum, urine and water.

The unique features of the kit are:

- Rapid extraction method for various samples with high recovery (>80%).

- High sensitivity (0.5 ng/g or ppb) and low detection limit (0.25 ng/g or ppb for shrimp).

- A quick ELISA assay (less than 2 hours regardless of number of samples).

- High reproducibility..

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The plate wells are coated with capture antibodies directed against anti-SQX antibodies. SQX standards or sample solution, SQX –HRP conjugate and anti-SQX antibodies are added. Free SQX and SQX-HRP Conjugate compete for the SQX antibody binding sites. At the same time, the anti-SQX antibodies are also bound by the immobilized capture antibodies. Any unbound SQX-HRP Conjugate is then removed in a washing step. The TMB Substrate is added to the wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Sulfaquinoxaline ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C*. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Sulfaquinoxaline Plate | 1x96-well Plate(8wellsx12 strips) | 2-8C |
| SQX Standards: Negative control (white cap tube) 0.5 ng/mL (yellow cap tube) 1.5 ng/mL (orange cap tube) 5.0 ng/mL (pink cap tube) 15 ng/mL (purple cap tube) 50 ng/mL (blue cap tube) 1000 ng/mL (spiking, optional, red cap tube) | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | 2-8C 2-8C |
| SQX Antibody | 6 mL | 2-8C |
| SQX-HRP Conjugate | 6 mL |  |
| 20X Wash Solution ** | 28 mL |  |
| Stop Buffer ** | 14 mL |  |
| TMB Substrate ** | 10 mL |  |
| Concentrate of Tissue Extraction Buffe | 60g |  |
| Tissue Clean Up Mix | 12g |  |
| 10X Sample Extraction Buffer E | 25 mL |  |

* If you are not planning to use the kit for over 1 month, store SQX Antibody and SQX-HRP Conjugate at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Fish/Shrimp/Meat/Liver/Kidney | 0.25 |
| Honey | 1.0 |
| Milk | 2.5 |
| Serum | 2.5 |
| Urine | 2.5 |
| Water | 0.25 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Sulfaquinoxaline | 100 |
| Sulfamethoxypyridazine | 70 |
| Sulfadiazine | 4 |
| Sulfamoxole | 1 |
| Sulfisoxazole | 0.5 |
| Sulfathiazole | <0.1 |
| Sulfamerazine | <0.1 |
| Sulfaquanidine | <0.1 |
| Sulfadoxine | <0.1 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- N-hexane

- Acetonitrile

- Ethyl Acetate

### Warnings and Precautions

- The standards contain sulfaquinoxaline. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES, HRP CONJUGATE AND PLATES ARE KIT-AND LOT-SPECIFIC.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature 20°–25°C (68°–77°F). while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

1.Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 10X Sample Extraction Buffer E with 9 volumes of distilled water.

The recommended dilution factors for different type of samples shown below are minimum ones in order to achieve the lowest detection limit. The dilution factor can be tailored to meet larger sample positive cutoff value.

### Fish/Shrimp/Meat/Liver/Kidney

- Homogenize a reasonable amount of sample with a suitable mixer.

- Weigh out 3 g of the homogenized sample and mix with 6 mL of ethyl acetate, vortex 3 min at max speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature 20°–25°C (68°–77°F).

- Transfer 4 mL of the ethyl acetate supernatant (corresponding to 2 g of the original sample) into a new vial and use a rotary evaporator to dry the sample in a 50-60°C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50-60°C water bath.

- Dissolve the dried residue in 2 mL of n-hexane.

- Add 1 mL of 1X Sample Extraction Buffer and mix by vortexing at maximum speed for 1 minute.

- Centrifuge for 10 minutes at 4,000 x g at room temp. Discard the upper hexane layer.

- Use 50 L of the lower aqueous layer per well for the assay. In case emulsion happened, remove the upper n-hexane layer and incubate the lower aqueous layer in water bath for 3 min at 80-95°C.

Note: Dilution factor: 0.5. If needed, the sample obtained from step 7 can be diluted with 1X Sample Extraction Buffer.

### Honey

- Add 1 mL of 0.5 M HCl in 1 g of honey sample, incubate 30min at 37°C.

- Add 2.5 mL of 0.2 M NaOH to the sample and mix well.

- Add 4 mL of ethyl acetate, shake for 10 min, centrifuge at 4000 rpm at room temperature for 10 min.

- Take 2 mL supernatant, blow to dry with nitrogen at 50°C, add 1 mL of 1X Sample Extraction Buffer, vortex and dissolve the sample.

- Use 50 L of the sample per well for the assay.

Note: Dilution factor: 2. If needed, the sample obtained from step 4 can be diluted with 1X Sample Extraction Buffer.

### Milk

- For fat-free milk, dilute the milk sample with 1X Sample Extraction Buffer (1:4) (e. g. 20 L of sample + 80L of buffer). Take 50L of the diluted sample per well for the assay.

- For the regular milk with fat, centrifuge the milk sample at 4,000 x g for 5 minutes, discard the upper fat layer. Dilute the sample with 1X Sample Extraction Buffer (1:4) (e. g. 20 L of milk sample + 80L of buffer). Take 50 L of the diluted sample per well for the assay.

Note: Dilution factor: 5. If needed, the sample can be further diluted with 1X Sample Extraction Buffer.

### Serum

- Centrifuge 0.5 mL of the serum sample at 4,000 x g for 5 minutes.

- Take out 20 L of the supernatant, add 80 L of 1X Sample Extraction Buffer, and mix well.

- Use 50 L per well for the assay.

Note: Dilution factor: 5. If needed, the sample can be further diluted with 1X Sample Extraction Buffer.

### Urine

- Centrifuge 0.5 mL of the serum sample at 4,000 x g for 5 minutes.

- Take out 20 L of the supernatant, add 80 L of 1X Sample Extraction Buffer, and mix well.

- Use 50 L per well for the assay.

Note: Dilution factor: 5. If needed, the sample can be further diluted with 1X Sample Extraction Buffer.

### Water

- Take out 5 mL of water sample. If the sample has been preserved with monochloroacetic acid or other acids, the sample should be neutralized with 6N NaOH prior to extraction. Centrifuge the water sample at 4,000 x g for 5 minutes.

- Take out 3 mL of the supernatant, add 0.5 mL of 2 M NaCI, and mix with 6 mL of ethyl acetate, vortex 3 min at max speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 4 mL of the ethyl acetate supernatant (corresponding to 2 mL of the original water sample) into a new vial and use a rotary evaporator to dry the sample in a 50-60C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50-60C water bath.

- Add 1 mL of 1X Sample Extraction Buffer to the dried sample and mix by vortexing at maximum speed for 1 minute.

- Use 50 L of the sample per well for the assay.

Note: Dilution factor: 0.5 If needed, the sample obtained from step 5 can be diluted with 1X Sample Extraction Buffer.

SQX ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Sulfaquinoxaline Antibody Solution | 50 L | 1.2 mL |
| Sulfaquinoxaline -HRP Conjugate | 50 L | 1.2 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

- Add 50 L of each SMX Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration). Add 50 L of each sample in duplicate into different sample wells.

- Add 50L of Sulfaquinoxaline -HRP Conjugate to each well

- Add 50L of Sulfaquinoxaline Antibody Solution to each well and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 1 hour at room temperature (20 – 25C / 68 – 77F).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating ( Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Sulfaquinoxaline Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical SQX standard curve. The sample detection and quantification limit are calculated as below.

Sample detection limit = (0.5 ng/g) x (dilution factor)

Sample quantification limit = (1.5 ng/g) x (dilution factor)

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com   reagenllc@gmail.com

Web.: www.reagen.us