REAGEN™

Crystal Violet/LCV ELISA Test Kit

RND99044

### Product Description

REAGEN™ Crystal Violet/LCV ELISA Test Kit (patent-pending) is designed for routine screening for Crystal Violet/Leuco crystal Violet in fish, shrimp, fish/shrimp feed, fish oil and water samples of fish farms or fishponds .The unique features of the kit are:

- High recovery (80 - 95%), and rapid extraction method for fish/shrimp or water samples.

- Converting leucocrystal Violet (LCV) to Crystal Violet before detection. Chromatic Crystal Violet can be rapidly bioconverted into colorless LCV in fish (~80%). The special Oxidant Solution in the kit can convert LCV back to the Crystal Violet form, which can be detected by the kit.

- Providing a 96-well high throughput quantitative detection method. The detection limit is 0.1 ng/g or ppb for fish/shrimp samples.

- The prepared samples can be directly used for liquid chromatography confirmation.

- High reproducibility.

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The Crystal Violet (CV) antibody has been coated in the plate wells. During the analysis, sample is added along with the CV-horseradish peroxidase (CV-HRP) conjugate. If the CV residue is present in the sample, it will compete for the CV antibody, thereby preventing the CV-HRP from binding to the antibody attached to the well. The resulting color intensity, after addition of the HRP substrate (TMB), has an inverse relationship with the CV residue concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Crystal Violet/LCV ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). The kit components should be stored at 2-8C*. Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Crystal Violet-coated Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8C |
| Crystal Violet Standards: Negative control (white cap tube) 0.05 ng/mL (yellow cap tube) 0.15 ng/mL (orange cap tube) 0.5 ng/mL (pink cap tube) 1.5 ng/mL (purple cap tube) 4.5 ng/mL (blue cap tube) 100 ng/mL (spiking, optional, red cap tube) | 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL | 2-8C |
| CV-HRP Conjugate | 8 mL | 2-8C |
| 20X Wash Solution ** | 28 mL |  |
| Stop Buffer ** | 14 mL |  |
| TMB Substrate ** | 12 mL |  |
| Concentrate of Sample Extraction Buffer A ** | 20 g (containing Reagent I and II) |  |
| 10X Sample Extraction Buffer B ** | 10 mL |  |
| 10X Sample Extraction Buffer C ** | 25 mL |  |
| 10X Oxidant Solution ** | 1.8 mL |  |

* 	If you are not planning to use the kit for over 3 months, store CV-HRP Conjugate and Crystal Violet standards at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Fish/Shrimp | 0.1 |
| Fish/Shrimp Feed | 0.2 |
| Fish Oil | 0.1 |
| Water Sample* | 0.07 for Fresh Water; 0.05 for Seawater |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Crystal Violet | 100 |
| Leucocrystal Violet | 100 (after oxidation) |
| Malachite Green | 82 |
| Leucomalachite Green | <0.01 |
| Diethylstilbestrol | <0.01 |
| Sulfumethazine | <0.01 |
| Chloramphenicol | <0.01 |
| Furazolidone | <0.01 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Rotary evaporator or nitrogen gas

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Dichloromethane

- Acetonitrile

- n-Hexane

### Warnings and Precautions

- The standards contain Crystal Violet. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. CV-HRP CONJUGATE AND PLATES ARE KIT-AND LOT-SPECIFIC.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for

no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X Oxidant Solution

Mix 1 volume of 10X Oxidant Solution with 9 volumes of acetonitrile.

- Preparation of 1X Sample Extraction Buffer A

Carefully open Sample Extraction Buffer A Concentrate bag and take out Reagent I and II and mix them in a bottle. Add 100 mL of distilled water to dissolve the reagents. The buffer can be stored at room temperature (20 – 25C / 68 – 77F) ( Since the reagents are corrosive please wear protection gloves when preparing the buffer).

- Preparation of 1X Sample Extraction Buffer B

Mix 1 volume of 10X Sample Extraction Buffer B with 9 volumes of distilled water.

- Preparation of 1X Sample Extraction Buffer C

Mix 1 volume of 10X Sample Extraction Buffer C with 9 volumes of distilled water.

### Fish/Shrimp

- Blend boneless fish/shrimp tissue in a blender/homogenizer until the sample has consistency of fine paste. Store the sample at -20ºC or in a freezer till analysis.

- Take 2.0 g of homogenized fish/shrimp sample, add 0.6 mL of 1X Sample Extraction Buffer A, 0.4 mL of 1X Sample Extraction Buffer B and 6.5 mL of acetonitrile. Vortex for 3 minutes at maximum speed or rotorack for 20 minutes (e.g. VWR Tube Rotator). Add 3.0 mL of water and 2.0 mL of dichloromethane to the sample, mix by gently inverting the tube for 1 min.

- Centrifuge the sample for 5 minutes at 4,000 x g, take out 4 mL of the upper organic solvent layer ( The total upper layer is ~7 mL, avoid the lower aqueous layer!) to a new tube containing 0.1 mL of 1X Oxidant Solution. Keep the sample at room temperature (20 – 25C / 68 – 77F) for 15 minutes and then store the sample at -20ºC or in a freezer for 1 hour (could be 1-3 hours).

- Take out the cold sample and immediately centrifuge the sample for 5 minutes at 4,000 x g. Take out 3 mL of the upper solvent layer and use a rotary evaporator to dry the sample in a 50 - 60C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50 - 60C water bath.

- Add 2 mL of n-hexane to dissolve the sample and then add 1.6 mL of the solution [1X Sample Extraction Buffer C: acetonitrile, 4:1, V/V], mixing the sample by inverting the tube for 1 minute.

- Leave the tube open and heat the sample at 85C for 3 minutes ( This step can be omitted if the lower aqueous layer is enough for the ELISA assay).

- Centrifuge the sample for 10 minutes at 4,000 x g.

- Discard the upper organic layer. Use 100 L of the lower aqueous layer for the assay.

Note: Dilution factor: 2. (2 g fish x 3/7.25 =~ 0.8 g in 1.6 mL of the solution [1X Sample Extraction Buffer C]. The prepared sample can be tested within 24 hours.

### Fish/Shrimp Feed

- To a reasonable amount of feed sample, add the equal amount of distilled water (e.g. 10 g of sample + 10 mL of water); homogenize the sample with a suitable mixer. Store the sample at -20ºC or in a freezer till analysis.

- Take 2.0 g of the homogenized sample, add 0.6 mL of 1X Sample Extraction Buffer A, 0.4 mL of 1X Sample Extraction Buffer B and 6.5 mL of acetonitrile. Vortex for 3 minutes at maximum speed or rotorack for 20 minutes. Add 3.0 mL of water and 2.0 mL of dichloromethane to the sample, mix by gently inverting the tube for 1 min.

- Centrifuge the sample for 5 minutes at 4,000 x g, take out 4 mL of the upper organic solvent layer ( The total upper layer is ~7 mL, avoid the lower aqueous layer!) to a new tube containing 0.1 mL of 1X Oxidant Solution. Keep the sample at room temperature (20 – 25C / 68 – 77F) for 15 minutes and then store the sample at -20ºC or in a freezer for 1 hour (could be 1-3 hours).

- Take out the cold sample and immediately centrifuge the sample for 5 minutes at 4,000 x g. Take out 3 mL of the upper solvent layer and use a rotary evaporator to dry the sample in a 50 - 60C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50 - 60C water bath.

- Add 2 mL of n-hexane to dissolve the sample and then add 1.6 mL of the solution [1X Sample Extraction Buffer C: acetonitrile, 4:1, V/V], mixing the sample by inverting the tube for 1 minute.

- Centrifuge the sample for 10 minutes at 4,000 x g.

- Discard the upper organic layer. Use 100 L of the lower aqueous layer for the assay.

Note: Dilution factor: 4.

### Fish Oil

- Take 2.0 g of fish oil sample, add 0.6 mL of 1X Sample Extraction Buffer A, 0.4 mL of 1X Sample Extraction Buffer B and 6.5 mL of acetonitrile. Vortex for 3 minutes at maximum speed or rotorack for 20 minutes (e.g. VWR Tube Rotator). Add 3.0 mL of water and 2.0 mL of dichloromethane to the sample, mix by gently inverting the tube for 1 min.

- Centrifuge the sample for 5 minutes at 4,000 x g, take out 4 mL of the upper organic solvent layer ( The total upper layer is ~7 mL, avoid the lower aqueous layer!) to a new tube containing 0.1 mL of 1X Oxidant Solution. Keep the sample at room temperature (20 – 25C / 68 – 77F) for 15 minutes and then store the sample at -20ºC or in a freezer for 1 hour (could be 1-3 hours).

- Take out the cold sample and immediately centrifuge the sample for 5 minutes at 4,000 x g. Take out 3 mL of the upper solvent layer and use a rotary evaporator to dry the sample in a 50 - 60C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50 - 60C water bath.

- Add 2 mL of n-hexane to dissolve the sample and then add 1.6 mL of the solution [1X Sample Extraction Buffer C: acetonitrile, 4:1, V/V], mixing the sample by inverting the tube for 1 minute.

- Leave the tube open and heat the sample at 85C for 3 minutes ( This step can be omitted if the lower aqueous layer is enough for the ELISA assay).

- Centrifuge the sample for 10 minutes at 4,000 x g.

- Discard the upper organic layer. Use 100 L of the lower aqueous layer for the assay.

Note: Dilution factor: 2. The prepared sample can be tested within 24 hours.

### Water Sample of Fish Farms or Fishponds

-- Fresh Water Sample --

- Centrifuge the water sample at 4,000 x g for 5 minutes.

- Take 710 L of the supernatant, add 10 L of 1X Sample Extraction Buffer B, mix well and leave at room temperature for 20 min.

- Add 80 L of 10X Sample Extraction Buffer C and 200 L of acetonitrile to the tube, mix well. The sample is ready for ELISA assay.

Note: Dilution factor: 1.41

-- Seawater Sample --

- Centrifuge the seawater sample at 4,000 x g for 5 minutes.

- Take out 1 mL of seawater sample, add 0.5 mL of dichloromethane, vortex 1 min at max speed.

- Centrifuge the sample for 5 minutes at 4,000 x g, remove the upper water layer, take 0.25 mL of lower organic solvent layer and dry it by blowing nitrogen gas in a 60-70C water bath.

- Add 0.5 mL of the solution [1X Sample Extraction Buffer C: acetonitrile, 4:1, V/V], vortex for 1 minute. Leave the sample at room temperature for 20 min.

- Use 100 L of the sample for the assay.

Note: Dilution factor: 1.

CRYSTAL VIOLET/LCV ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Buffer concentrate with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| CV-HRP Conjugate | 50mL | 1.2 mL |
| 1X Wash Solution | 1.0 mL | 24 mL |
| Stop Buffer | 100 mL | 2.4 mL |
| TMB Substrate | 100 mL | 2.4 mL |

- Add 100 mL of each CV Standards in duplicate into different wells (F Add standards to plate only in the order from low concentration to high concentration).

- Add 100 mL of each sample in duplicate into different sample wells.

- Add 50 mL of CV-HRP Conjugate and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 1 hour at room temperature (20 – 25°C / 68 – 77°F) in the dark.

- Wash the plate 3 times with 250 mL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (F Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 mL of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. Incubate for 20 minutes at room temperature (20 – 25°C / 68 – 77°F) in the dark. (Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubation, add 100 mL of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength (F Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Crystal Violet Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical Crystal Violet standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com    reagenllc@gmail.com

Web.:www.reagen.us

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-671-1789         FAX.:856-858-5434

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web. www.reagen.us