REAGEN™

Melamine ELISA Test Kit

RND99010

Product Description

REAGEN™ Melamine ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of melamine in egg,milk, milk powder, infant formula, meat, fish, feed, chocolate yogurt, butter milk ,sour milk and other food  samples. The unique features of the kit are:

- High recovery (>80%),rapid(10-40 minutes),and cost-effective extraction methods.

- High sensitivity(5ng/g or ppb)and low detection limit for liquid milk (0.2ppm).

- A quick ELISA assay(less than 2 hours regardless of number of samples).

- High reproducibility.

Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target toxin and enzyme-conjugate. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The enzyme-conjugate, will bind with the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

Kit Contents, Storage and Shelf Life

REAGEN™ Melamine ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8°C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Melamine Coated Plate | 1 x 96-well Plate (8 wells x 12 strips) | 2-8°C |
| Melamine Standards: Negative control (white cap tube) 5 ng/mL (yellow cap tube) 20 ng/mL (orange cap tube) 50 ng/mL (pink cap tube) 100 ng/mL (purple cap tube) 200 ng/mL (blue cap tube) 100 µg/mL (spiking, optional, red cap tube) | 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL | 2-8°C * |
| Melamine Antibody | 8 mL | 2-8°C * |
| 1X HRP Conjugate | 8 mL | 2-8°C * |
| 20X Wash Solution | 28 mL | 2-8°C |
| Stop Buffer | 20 mL | 2-8°C |
| TMB Substrate | 12 mL | 2-8°C |
| 10 X Melamine Sample Extraction Buffer | 20 mL X 2 | 2-8°C |
| 1 X Melamine Clean Up Buffer | 20 mL | 2-8°C |
| 1X Melamine Sample Balance Buffer | 20 mL | 2-8°C |
| 10X Melamine Sample Dilution Buffer | 14 mL | 2-8°C |
| 10X milk Extraction Buffer(optional) | 5 mL | 2-8°C |

*If you are not planning to use the kit for over 1 month, store  Melamine     Standards, Melamine Antibody and  1X HRP Conjugate at -20°C or in a freezer.

Sensitivity (Detection Limit)

| Sample Type | Detection Limit (g/g or ppb) |
| --- | --- |
| Egg | 0.055ppm |
| Feed | 0.5ppm |
| Meat /fish | 0.055ppm |
| Milk | 0.2ppm |
| Milk powder/infant formula | 0.4ppm |
| Chocolate | 0.4ppm |
| Yogurt/butter milk /sour milk | 0.4ppm |

Specificity (Cross-Reactivity)

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Melamine | 100 |
| Cyanurate | 15 |
| Ammeline | 5 |
| Ammelide | <1 |

Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Genie Vortex mixer from VWR)

- 10, 20, 100 and 1000 µL pipettes

- Multi-channel pipette: 50-300 µL (Optional)

Warnings and Precautions

- The standards contain melamine. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within  their  expiration  dates. ANTIBODIES  AND PLATES  ARE  KIT-AND LOT-SPECIFIC.  Make sure that the HRP Conjugate and Diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays  in  direct sunlight,  as this may cause excessive  heat and evaporation. Cold  bench tops should be avoided by placing several layers of paper towel or some other insulate material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents  into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards  first and then your samples.

- Add standards  to plate only in the order from low concentration to high concentration, as this will minimize the risk of compromising the standard curve.

- Always  refrigerate  plates  in  sealed  bags  with  a  desiccant  to  maintain  stability.  Prevent condensation from forming on plates by allowing them equilibrate to room temperature  (20 – 25°C /68 – 77°F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C  if  they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25°C / 68 – 77°F) or in a refrigerator before use.

1. Preparation of 1X Melamine Sample Extraction Buffer

Mix 1 volume of the 10 X Melamine Sample Extraction Buffer with 9 volumes of distilled water.

2. Preparation of 1X Melamine Sample Dilution Buffer

Mix 1 volume of the 10 X Melamine Sample Dilution Buffer with 9 volumes of distilled water.

- Preparation of 1X Milk Extraction Buffer

Mix 1 volume of 10X Milk Extraction Buffer with 9 volumes of distilled water.

Egg

- Weigh out 1 g of homogenized egg sample, add 4 mL 1X Melamine Sample Extraction Buffer to all the samples. Vortex about 30 seconds to get a homogeneous solution.

- Add 250 µL of Melamine Clean Up Buffer to each tube, vortex at maximum speed for 3 minutes manually or vortex in a multi-vortexer for 10 minutes.

- Centrifuge at 4,000 X g for 10 minutes.

- Transfer 300 µL of supernatant to another tube, dilute with 300 µL of 1X Melamine Sample Balance Buffer and 1 mL of hexane.

- Vortex for 60 seconds. Centrifuge at 4,000 x g for 5 minutes.

- Remove the top hexane layer completely. Use 50 µL of the solution from the bottom layer per well for the assay (Try to avoid the middle cake-like layer).

Note: Dilution factor: 11. (fit for 0.055 ppm to 2 ppm spiking).

Feed

1.	Grind and mix a representative sample (according to accepted sampling techniques).

2.	Weigh out 0.5 g of the ground sample and place into a suitable container.

3.	Add 5 mL of 1 X Sample Extraction Buffer to all the samples,shake the samples for 5 seconds. Then add 250 µL Melamine Clean Up Buffer to each sample, vortex at maximum speed for 3 minutes or in a multi-vortexer for 10 minutes.

4.	Centrifuge at 4000 xg for 10 minutes.

5.	Transfer 250 µL of clear supernatant to another tube, dilute with 250 µL of 1 X Sample Balance Buffer .

6.	Vortex for 15 seconds. Centrifuge at 4,000 x g for 2 minutes.

7.	Use 50 µL of the sample per well for the assay.

Note: Dilution factor: 22.

To detect a higher concentration of melamine in the sample, the dilution factor can be  increased by adding 1X Melamine Sample Dilution Buffer in Step 5.  For example, if 500 µL of 1X Melamine Sample Dilution Buffer is  the dilution factor is 44.

Meat/Fish

- Weigh 1 g of  meat/fish sample.

2.       Add 4 mL 1X Melamine Sample Extraction Buffer to all the samples. Shake the samples for 5 seconds. Then add 250  µL of Melamine Clean Up Buffer to each sample, vortex at maximum speed for 3 minutes manually or vortex in a multi-vortexer for 10 minutes.

3.       Centrifuge at 4,000 X g for 10 minutes.

4.       Transfer 250 µL of clear supernatant to another tube, dilute with 250 µL of 1X Melamine Sample Balance Buffer.

5.       Vortex for 15 seconds. Centrifuge at 4,000 x g for 2 minutes.

6.       Use 50 µL of the sample per well for the assay.

Note: Dilution factor: 11. (fit for 0.2 ppm to 2 ppm spiking).

To detect a higher concentration of melamine in the sample, the dilution factor can be increased by adding 1X Melamine Sample Dilution Buffer in Step 4.  For example, if 500 µL of 1X Melamine Sample Dilution Buffer is added, the dilution factor is 22 (fit for 0.4 ppm to 4 ppm spiking).

Milk

1.	Take 1 mL of milk  sample,  add 4ml of Melamine  Sample Extraction Buffer to all the

samples. Shake the sample for 5 seconds. Then add 250 µL of Melamine Clean Up Buffer to each tube,  vortex at  maximum  speed for  3 minuutes seconds or  shake

for  10 minutes  in  a multi-vortexer or shaker.

2.	Centrifuge at 4,000 x g for 10 minutes.

3.	Transfer 250 µL of clear supernatant to another tube,dilute with 250 µL of

1X Melamine Sample Balance Buffer

4.     Vortex for 15 seconds. Centrifuge at 4,000 x g for 2 minutes.

5.      Use 50 µL of the sample per well for the assay.

Note: Dilution factor:11  (fit for 0.2 ppm to 2.0 ppm spiking).

To detect a higher concentration of melamine in the sample, the dilution factor can be increased by adding 1X Melamine Sample Dilution Buffer in Step 4.  For example, if 500 µL of 1X Melamine Sample Dilution Buffer is added, the ilution factor is 22 (fit for 0.4 ppm to4.0 ppm spiking).

Milk Powder/Infant Formula

1.	Take 1 g of milk powder/Infant formula, add 4 mL of  1X Melamine Sample Extraction Buffer. Vortex about 30 seconds to get a homogeneous solution.

2.	Add 250 µL of Melamine Clean Up Buffer to each sample, vortex at maximum speed for 3 minutes,  manually or vortex in a multi-vortexer for 10 minutes.

3.	Centrifuge at 4,000 x g for 10 minutes.

4.	 Transfer 100 µL of clear supernatant to another tube, dilute with 100 µL of 1X Melamine Sample Balance Buffer.

5.     Vortex for 15 seconds. Centrifuge at 4,000 x g for 2 minutes.

6.	Use 50 µL of the sample per well for the assay.

Note: Dilution factor: 11. (fit for 0.2 ppm to 2 ppm spiking).

To detect a higher concentration of melamine in the sample, the dilution factor can be increased by adding 1X Melamine Sample Dilution Buffer in Step 4.  For example, if 800 µL of1X Melamine Sample Dilution Buffer is added, the dilution factor is 55 (fit for 0.6  ppm to10 ppm spiking ).

Chocolate

1.  Weigh 1 g of Chocolate into a GLASS TUBE, add 5 mL 1X Melamine Sample Extraction Buffer.  Incubate in a  50 °C water bath for 2-3 minutes, vortex about 30 seconds at maximum speed to get a homogeneous suspension.

2.  Add 250 µL of Melamine Clean Up Buffer, vortex at maximum speed for 3 minutes manually or vortex in a multi-vortexer for 10 minutes.

3.  Centrifuge at 4,000 X g for 10 minutes.

4.  Transfer 250 µL of clear supernatant to another tube, dilute with 250 µL of 1X Melamine Sample Balance Buffer.

5.  Vortex for 15 seconds. Centrifuge at 4,000 x g for 2 minutes.

6.  Use 50 µL of the sample per well for the assay.

Note: Dilution factor: 11. (fit for 0.2 ppm to 2 ppm spiking).

To detect a higher concentration of melamine in the sample, the dilution factor can be increased by adding 1X Melamine Sample Dilution Buffer in Step 4.  For example, if 500 µL of 1X Melamine Sample Dilution Buffer is added, the dilution factor is 22 (fit for 0.4 ppm to 4 ppm spiking).

Yogurt/Buttermilk/Sour Milk

- Weigh 0.5 g of the sample, add 0.5 mL 1X Milk Extraction Buffer to all the samples. Vortex at maximum speed for 1 minute manually or vortex in a multi-vortexer for 5 minutes.

- Centrifuge at 4,000 X g for 10 minutes.

- Transfer 200 µL of clear supernatant to another tube, dilute with 800 µL of 1X Melamine Sample Dilution Buffer.

- Mix Well. Use 50 µL of the sample per well for the assay.

Note: Dilution factor: 10. (fit for 0.2 ppm to 2 ppm spiking). To detect a higher concentration of melamine in the sample, the dilution factor can be increased to 50 by adding 960 µL1X Melamine Sample Dilution Buffer to 40 µL of clear supernatant in Step 3 (fit for 1.0 ppm to 10 ppm spiking).

MELAMINE ELISA TEST KIT PROTOCOL

Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at20 – 25°C / 68 – 77°F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.    All reagents should be mixed by gently inverting or swirling prior to use. Prepare only the volumes that are needed for the number of wells being run. Do  not return the  reagents  to the  original stock  tubes/bottles. It  is recommended  that disposable reservoirs be used when handling reagents to minimize the risk of contamination.

1.  Preparation of 1X Wash Solution

Mix 1 volume of the 20 X Wash Solution  with 19 volumes of distilled water.

ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Melamine Antibody | 50 µL | 1.2 mL |
| 1X HRP Conjugate | 50 µL | 1.2 mL |
| 1X Wash Solution | 1.0 mL | 24 mL |
| Stop Buffer | 100 µL | 2.4 mL |
| TMB Substrate | 100 µL | 2.4 mL |

1.	Add 50 µL  of each Melamine  Standards in  duplicate  into different  wells  (    Add standards to plate only in the order from low concentration to high concentration).

2.	Add 50 µL of each sample in duplicate into different sample wells.

3.	Add 50 µL of 1X HRP Conjugate to each well.

3.	Add 50 µL of Melamine Antibody and mix well by gently rocking the plate manually for 1 minute. (Add the HRP Conjugate first and then add Melamine Antibody)

5.	Incubate the plate for 30 minutes at room temperature (20 – 25°C / 68 – 77°F).

6.	Wash the plate 3 times with 250  µL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (    Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

7.	Add 100 µL of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (  Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

8.	After incubating for 20 minutes at room temperature (20 – 25°C / 68 – 77°F), add 100 µL of  Stop Buffer to stop the enzyme reaction.

9.	Read the plate as soon as possible following the addition of Stop  Buffer on a plate reader with 450 nm wavelength (  Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

Melamine Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical Melamine  standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250        FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us