REAGEN™

Fumonisin Toxin ELISA Test Kit

RNM98009

### Product Description

REAGEN™Fumonisin ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Fumonisin in cereals, dried meat/fish, feed, milk, seed and oils.

The unique features of the kit are:

- High recovery (80-120%),rapid(10-20minutes), cost-effective extraction methods.

- High sensitivity (1 ng/g or ppb).

- High reproducibility.

- A quick ELISA assay (less than 1 hours regardless of number of samples).

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The toxin of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target toxin and enzyme-conjugate. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the toxin attached to the well. The enzyme-conjugate, will bined with the primary antibody that is complexed to the toxin coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™Fumonisin ELISA Test Kit has the Fumonisinacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Fumonisin Antibody-Coated Plate | 1 x 96-well Plate (8 wells x 12 strips) | 2-8C |
| Fumonisin Standards: Negative control (white cap tube) 1 ng/mL (yellow cap tube) 5 ng/mL (orange cap tube) 10 ng/mL (pink cap tube) 25ng/mL (purple cap tube) 50 ng/mL (blue cap tube) 1000 ng/mL (spiking, red cap tube) | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | 2-8C 2-8C |
| Fumonisin Antibody #1 | 8ml | 2-8C |
| 1X HRP-Conjugated Antibody #2 | 8ml |  |
| 20X Wash Solution ** | 30 mL |  |
| Stop Buffer ** | 20 mL |  |
| TMB Substrate ** | 10 mL |  |

* 	If you are not planning to use the kit for over 3 months, store Fumonisin  Antibody #1 and 1X HRP-Conjugated Antibody #2 at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Meat/fish/sereal/feed/seed | 100 |
| Milk | 10 |
| Oils | 10 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Fumonisin B1 | 100.0 |
| Fumonisin B2 | 94 |
| Fumonisin B3 | 145 |
| T-2 | <0.1 |
| DON | <0.1 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Methanol

### Warnings and Precautions

- The standards contain Fumonisin . Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

### Dried meat/fiah/cereal/feed/seed

    - Grind and mix a representative sample(according to accepted sample techniques).

    - Weight out 5 g of the ground sample and place into a suitable container .

    - Add 25mL70% methanol and shake for 20 minutes with a shaker or multi-tube vortexer.

    - Centrifuger sample for 10 minutes at 4000 rpm.

    - Dilute 50L of the obtained supernatant with 0.95 ml of 35% methanol.

    - Use 50L of the diluted supernatant per well in the test.

Note: Dilution factor: 100.

### Milk

    - For fat-free milk, dilute the milk sample with 35% methanol in (1:9). (e.g. 20L of milk +1 80L of 35% methanolr). Take 50 L of the diluted sample per well for the assay.

    - For the regular milk with fat, centrifuge the milk sample at 4,000 x g for 5 minutes, discard the upper fat layer. Dilute the samples with 35% methanol (1:9). (e.g. 20L of milk +180L of 35% methanol). Take 50 L of the diluted sample per well for the assay.

Note: Dilution factor: 10.

### Vegatable/bean/peanut/olive oils

    - weign out 1.0 g of the oil sample,add 5.0 mL of 35% methanol in dH2O

    - vortex vigorously for 10 minutes in a multi-tube vortex or shaker,or vortex 3 minutes at maximum speed manually.

    - Transfer 200L of the subnatant diluted sample to a new tube and add 800L of 35% methanol in dH2O , vortex for 1 minutes .

    - Take 50 L of the diluted sample per well for the assay.

Note: Dilution factor: 25.

FUMONISIN ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Buffer concentrate with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Fumonisin Antibody #1 | 50L | 1.2mL |
| 1X HRP-Conjugated Antibody #2 | 50 L | 1.2 mL |
| 1X Wash Solution | 1.0 mL | 24mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

    - Add 50Lof each Fumonisin Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

    - Add 50L of each sample in duplicate into different sample wells.

    - Add 50 L of 1X Antibody #2 to each well, gently rocking the plate manually for 10s and then add 50 L Antibody #1 to each well.Mmix well by gently rocking the plate manually for 1 minute.

    - Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F).

    - Wash the plate 3 times with 300 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. ( Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

    - After incubation for 10-15 minutes in room temperature, add 100 L of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Fumonisin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical Fumonisin standard curve:

REAGEN LLC

Add.:102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us