REAGEN™

Aflatoxin M1 ELISA Test Kit

RNM98001

### Product Description

REAGEN™Aflatoxin M1 ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Aflatoxin M1 in milk,milk powder, soured milk.

The unique features of the kit are:

- No need to extract the sample. The sample can be directly used for the ELISA

- High sensitivity (0.05 ng/g or ppb) .

- A quick ELISA assay(30 minutes ELISA reaction).

- High reproducibility.

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The plate wells are coated with capture antibodies directed against anti-Aflatoxin M1 antibodies. Aflatoxin M1 standards or sample solution, Aflatoxin M1–HRP conjugate and anti-Aflatoxin M1 antibodies are added. Free Aflatoxin M1 and Aflatoxin M1-HRP Conjugate compete for the Aflatoxin M1 antibody binding sites. At the same time, the anti- Aflatoxin M1 antibodies are also bound by the immobilized capture antibodies. Any unbound Aflatoxin M1-HRP Conjugate is then removed in a washing step. The TMB Substrate is added to the wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Aflatoxin M1  ELISA Test Kit has the Aflatoxin M1capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Aflatoxin M1 Antibody-Coated Plate | 1 x 96-well Plate(8 wells x 12 strips) | 2-8C |
| Aflatoxin M1 Standards: Negative control (white cap tube) 0.05ng/mL (yellow cap tube) 0.15ng/mL (orange cap tube) 0.3 ng/mL (pink cap tube) 0.9ng/mL (purple cap tube) 2.7ng/mL (blue cap tube) 10 ng/mL (spiking, red cap tube) | 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL | 2-8C |
| Aflatoxin M1 Antibody #1 | 6ml | 2-8C * |
| HRP-Conjugated Antibody #2 | 6 mL | 2-8C |
| 20X Wash Solution | 28 mL | 2-8C |
| Stop Buffer | 14mL | 2-8C |
| TMB Substrate | 12 mL | 2-8C |
| 10x PBS | 10 mL | 2-8C |

- If you are not planning to use the kit for over 3 months, store Aflatoxin M1  Antibody #1 and 1X HRP-Conjugated Antibody #2 at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Milk | 0.05 |
| Milk powder | 0.05 (corresponding to the dissolved liquid milk; 1:10) |
| Soured Milk | 0.1 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Aflatoxin M1 | 100.0 |
| Aflatoxin B1 | 36 |
| Aflatoxin B2 | 25 |
| Aflatoxin M2 | 19 |
| Aflatoxin G1 | 17 |
| Aflatoxin G2 | 3 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

### Warnings and Precautions

- The standards contain Aflatoxin M1 . Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using distilled or deionized water since water quality is important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

  - Preparation of 1X PBS:

Mix 1 volume of 10x PBS with 9 volumes of distilled water.

### Milk

In general, the raw milk and homogenized milk can be used directly for the assay. However, if high background is obtained with high fat milk, centrifuge the milk for 5 minutes, discard the upper fatty layer and use the remaining sample for the assay.

Note: Dilution factor: 1.0

### Milk powder

Weigh out 1 g of milk powder in a vial, add distilled water to 10 mL, dissolve sample by shaking. If high background is obtained for the fatty milk powder, centrifuge the dissolved milk for 5 minutes, discard the upper fatty layer and use the remaining sample for the assay.

Note: Dilution factor: 1.0 (corresponding to the dissolved liquid milk).

### Soured Milk

- Take out 0.5 mL of the sample into a vial, add 0.5 mL of 1 X PBS, vortex for 3 minutes at maximum speed.

- Take 50 L of the supernatant (avoid the top fat layer!) for the assay.

Note: Dilution factor: 2.

AFLATOXIN M1 ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions”

section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock

tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Buffer concentrate with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Aflatoxin M1 Antibody #1 | 50L | 1.2mL |
| HRP-Conjugated Antibody #2 | 50L | 1.2mL |
| 1X Wash Solution | 1.0 mL | 24mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

    - Add 50Lof each Aflatoxin M1 Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

    - Add 50L of each sample in duplicate into different sample wells.

    - Add 50L of HRP-Conjugated Antibody #2 to each well

    - Add 50L of Antibody #1 to each well and immediately following mixing in the well by pipetting up and down once.

    - Incubate the plate for 15minutes at room temperature (20 – 25C / 68 – 77F).

    - Wash the plate 4 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. ( Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

    - After incubation for 15 minutes in room temperature, add 100 L of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Aflatoxin M1 Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical Aflatoxin M1 standard curve

REAGEN LLC

Add.:102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250       FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us