REAGEN™

Okadaic Acid（DSP）ELISA Test Kit

RNA97008

### Product Description

REAGEN™Okadaic Acid (DSP) ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Okadaic Acid (DSP) in mussels and water samples.

The unique features of the kit are:

- High recovery (80-105%), rapid (10-30 minutes), and cost-effective extraction methods

- High sensitivity (0.3 ng/g or ppb) and low detection limit (30 ppb for mussel samples and 6 ppb for water samples).

- High reproducibility.

- A quick ELISA assay(less than 1.5 hours regardless of number of samples).

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The second antibody against Anti-Okadaic Acid Antibody was coated in the plate wells. During the analysis, sample is added along with the okadaic acid-horseradish peroxidase conjugate (Okadaic Acid-HRP Conjugate) and Anti- Okadaic Acid Antibody. If the okadaic acid residue is present in the sample, it will compete with Okadaic Acid-HRP Conjugate for the Anti-Okadaic Acid Antibody, thereby preventing the Okadaic Acid-HRP Conjugate from binding to the Anti-Okadaic Acid Antibody. The Anti-Okadaic Acid Antibody is then bound by the second antibody immobilized on the plate. After a washing step and addition of the TMB substrate solution, a color signal is produced. The color reaction is stopped after a specified time and the color is evaluated using an ELISA reader. The resulting color intensity has an inverse relationship with the okadaic acid residue concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Okadaic Acid (DSP) ELISA Test Kit has the Okadaic Acid (DSP) capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8℃ . The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| DSP Second Antibody-coated Plate | 1 x 96-well Plate(8 wells x 12 strips) | 2-8℃ |
| Okadaic Acid Standards: Negative control (white cap tube) 0.3ng/mL (yellow cap tube) 0.9ng/mL (orange cap tube) 2.7ng/mL (pink cap tube) 5.4ng/mL (purple cap tube) 16.2ng/mL (blue cap tube) | 0.8mL 0.8 mL 0.8 mL 0.8mL 0.8 mL 0.8mL | -20℃ |
| Anti-Okadaic Acid Antibody | 6 ml | 2-8℃ |
| Okadaic Acid-HRP Conjugate | 6 mL |  |
| 20X Wash Solution | 30 mL |  |
| Stop Buffer | 14mL |  |
| TMB Substrate | 12mL |  |
| 10X Sample Extraction Buffer | 25 mL |  |

If you are not planning to use the kit for over 1 months, store Okadaic Acid-HRP Conjugateand Anti- Okadaic Acid Antibody at -20℃ or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Mussels | 30 |
| Water | 6 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Okadaic Acid | 100.0 |
| DTX1 | 40 |
| DTX2 | 46 |
| Neosaxitoxin | < 0.01 |
| Saxitoxin | < 0.01 |
| YTX | < 0.01 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer

- Vortex mixer

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Methanol

### Warnings and Precautions

- The standards contain okadaic acid. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using distilled or deionized water since water quality is important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- 1.Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

- Preparation of 1X Sample Extraction Buffer/Methanol(90/10, V/V):

Mix 9 volumes of 1X Sample Extraction Buffer with 1 volume of methanol.

### Mussels

- Remove the mussel tissue samples from shells, wash with deionized water.

- Drain the excess liquid and homogenize the sample to a soupy texture.

- Weigh out 0.5 g of homogenized tissue, add 2 mL of 50% methanol (in water) and vortex for 5 minutes.

- Centrifuge tube for 10 minutes at 4,000 rpm.

- Transfer 0.5 mL of supernatant to a new tube, heat the sample at 75 °C for 5 minutes.

- Centrifuge 10 minutes at 4000 rpm.

- Transfer 50L of the clear supernatant to a new tube, add 950L of 1X Sample Extraction Buffer/Methanol (90/10, V/V), mix well. The sample is ready for the assay.

- Note: Dilution factor: 100 (detection range: 30-1620 ppb). Based on detection range request, the dilution factor can be adjusted higher or lower by changing the volume ratio of supernatant and 1X Sample Extraction Buffer/Methanol (90/10, V/V) in Step 7.

### Alternative Sample Preparation (use this method if hydrolysis is required)

- Remove the mussel tissue samples from shells, wash with deionized water.

- Drain the excess liquid and homogenize the sample to a soupy texture.

- Weigh out 0.5 g of homogenized tissue, add 2 mL of 80% methanol (in water) and vortex for 5 minutes.

- Centrifuge tube for 10 minutes at 4,000 rpm.

- Transfer 500L of the supernatant to a new tube, add 100L of 2.5 N NaOH. Cap the tube and vortex for 15-20 seconds.

- Incubate in a heating block set at 80℃ for 40 minutes.

- Allow to cool at room temperature. Add 100L of 2.5 N HCl, vortex for 15-20 seconds.

- Transfer 50L of the sample to a new tube, add 700L of 1X Sample Extraction Buffer/Methanol (90/10, V/V), mix well. The sample is ready for the assay.

- Note: Dilution factor: 105 (detection range: 32-1700 ppb). Based on detection range request, the dilution factor can be adjusted higher or lower by changing the volume ratio of sample and 1X Sample Extraction Buffer/Methanol (90/10,V/V) in Step 8.

### Water

- Take 50L of the water sample, add 950L of 1X Sample Extraction Buffer/Methanol(90/10, V/V).

- Take 50L of the diluted sample per well for the assay.

- Note: Dilution factor: 20 (detection range: 6-324 ppb). Based on detection range request, the dilution factor can be adjusted higher or lower by changing the volume ratio of water sample and 1X Sample Extraction Buffer/Methanol (90/10, V/V) in Step 1

OKADAIC ACID (DSP) ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions”

section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Buffer concentrate with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Okadaic Acid-HRP Conjugate | 50L | 1.2mL |
| Anti-Okadaic Acid Antibody | 50L | 1.2mL |
| 1X Wash Solution | 1.0mL | 24mL |
| Stop Buffer | 100 L | 2.4mL |
| TMB Substrate | 100 L | 2.4mL |

    - Add 50L of each Okadaic Acid Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

    - Add 50L of each sample in duplicate into different sample wells.

    - Add 50L of Okadaic Acid-HRP Conjugate to each well

    - Add 50L of Anti-Okadaic Acid Antibody to each well and mix well by gently rocking the plate manually for 1 minute.

    - Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F).

    - Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. ( Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

    - After incubation for 15 minutes in room temperature, add 100 L of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Okadaic Acid (DSP) Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical Okadaic Acid (DSP) standard curve