REAGEN™

Histamine Enzymatic Assay Kit

RNA97010

### Product Description

REAGEN™Histamine Enzymatic Assay Kit is a colorimetric enzymatic assay for the determination of histamine in fresh fish/seafood, fish flour, wine and milk.

The unique features of the kit are:

- Rapid, cost-effective aqueous extraction method with high recovery.

- New methanol-free extraction protocol

- High sensitivity and low detection limit (6 ppm for fresh fish/seafood).

- High reproducibility.

- A very rapid (10 minutes) and robust enzyme-based assay which does not require chemical derivatization or expensive instrumentation.

### Procedure Overview

The REAGEN™Histamine Enzymatic Assay Kit measures the chemical reduction of a dye molecule using a histamine-specific enzyme. The color change caused by reduction of the dye (measured by the absorbance at 450 nm) is related to the amount of histamine present in the sample.

### Kit Contents, Storage and Shelf Life

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Histamine-Microtiter Plate | 1x 96-well Plate (8wells x 12 strips) | -20C |
| Histamine standards : Negative control (white cap tube) 1.0 ppm (yellow cap tube) 2.0 ppm (orange cap tube) 4.0 ppm (pink cap tube) 8.0 ppm (purple cap tube) 12.0 ppm (blue cap tube) 1,000 ppm (spiking, optional, red cap tube) | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | -20C |
| 25X Sample Extraction Buffer | 25 mL | -20C |
| 100X Enrichment Solution | 25 mL | -20C |
| Master Mix ** | 6 mL | -20C |
| Enzyme Mix | 6 mL | -20C |

### Sensitivity

| Sample Type | Detection Limit (mg/kg or ppm) |
| --- | --- |
| Meat/Seafood | 6 |
| Fish flour | 20 |
| Milk | 40 |
| Wine | 40 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Histamine | 100 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Tissue Mixer

- Vortex Mixer

- Methanol

### Warnings and Precautions

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or different lots.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Add standards to plate only in the order from low concentration to high concentration, as this will minimize the risk of compromising the standard curve.

- After thawing the reagents, store reagents on ice or at 4°C (except for the extraction buffer, which can be stored at room temperature) to prevent degradation of the kit components.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.  Do not use glassware for extraction purposes. Histamine may adhere to glass which may affect test results.

### Meat/Seafood

MethodⅠ:

Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 25X Sample Extraction Buffer with 24 volumes of distilled water.

- Homogenize a reasonable amount of sample (e.g. 10 g). Transfer 1 g of the sample to a clean tube. Add 2 mL of methanol per gram of sample. Vortex mix for 2 minutes.

- Transfer 1.5 mL of the suspended homogenate to a microcentrifuge tube. Incubate the supernatant at 70°C for 5 minutes. Vortex mix for 5 seconds.

- Centrifuge for 5 minutes at 6,000 rpm (4,000 x g). Transfer 0.5 mL of supernatant to a new tube. ( Avoid transferring any precipitate or floating solid).

- Add an identical volume (0.5 mL) of 1X Sample Extraction Buffer to the sample.

Note: Dilution factor: 6. If possible, a known negative sample could be prepared in parallel; subtract the negative sample from the test sample results.

Method Ⅱ:

Preparation of 1X Enrichment Solution:

Mix 1 volumes of100X Enrichment Solution with 99 volumes of distilled water.

Preparation of 6X Sample Extraction Buffer:

Mix 6 volumes of 25X Sample Extraction Buffer with 19 volumes of distilled water.

- Homogenize a reasonable amount of sample (e.g. 20 g). Transfer 2 - 5 g (2 g is recommended) of the sample to a clean tube. Add 4 mL of 1X Enrichment Solution per gram of sample. (for example, add 8 ml of 1X Enrichment Solution to 2 g of homogenized fish). Vortex mix for 2 minutes.

- Transfer 1.2 - 1.5 mL of the suspended homogenate to a microcentrifuge tube. Heat the sample at 85°C for 10 minutes. Vortex mix for 5 seconds.

- Centrifuge for 5 minutes at 13,000 rpm (14,000 x g). Transfer 0.5 mL of supernatant to a new tube. ( Avoid transferring any precipitate or floating solid).

- Add 0.1 mL of 6X Sample Extraction Buffer to the sample. Mix briefly.

The samples are now ready to test in the assay.

Note: Dilution factor: 6. If possible, a known negative sample could be prepared in parallel; subtract the negative sample from the test sample results.

### Fish Flour

Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 25X Sample Extraction Buffer with 24 volumes of distilled water.

- Homogenize reasonable amount of the sample (e.g. 10 g). Take out 0.5 g of the sample to a tube; add 10 mL of 1X Sample Extraction Buffer. Vortex mix for 2 minutes.

- Centrifuge for 5 minutes at 6,000 rpm (4,000 x g). Transfer 1 mL of supernatant to a new tube. ( Avoid transferring any precipitate or floating solid).

- Incubate the supernatant at 75°C for 5 minutes. Vortex for 30 seconds.

- Centrifuge for 5 minutes at 6,000 rpm (4,000 x g). Transfer 0.4 mL of the supernatant to a new tube. ( Avoid transferring any precipitate or floating solid). Add an identical volume (0.4 mL) of methanol to the sample.

- Incubate the sample at 75°C for 5 minutes.

- Centrifuge for 5 minutes at 6,000 rpm (4,000 x g). Transfer 0.4 mL of the supernatant to a new tube.

Note: Dilution factor: 40. If possible, a known negative sample could be prepared in parallel; subtract the negative sample from the test sample results.

### Milk

Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 25X Sample Extraction Buffer with 24 volumes of distilled water.

- Centrifuge the milk sample for 5 minutes at 6,000 rpm (4,000 x g).

- Transfer the lower layer to a clean tube, avoiding contamination from the top layer of cream.

- Dilute the sample 1:40 in 1X Sample Extraction Buffer.

Note: Dilution factor: 40.

### Wine

Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 25X Sample Extraction Buffer with 24 volumes of distilled water.

1. Dilute the sample 1:40 in 1X Sample Extraction Buffer.

Note: Dilution factor: 40.

Histamine ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section . All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended. The thawed reagents can be stored at 4°C for as long as 2 weeks; for a longer time period storage, the reagents should be put in a freezer.

- Preparation of Reaction Mix

Thaw out Master Mix and Enzyme Mix and warm to room temperature (20 – 25°C / 68 – 77°F). Invert tube three times to mix thoroughly. In a clean tube, mix Master Mix and Enzyme Mix (1:1, V/V). Use the following table to calculate the required volume for the Reaction Mix:

| Reaction Mix |  | Volume per Reaction | 24 Reactions |
| --- | --- | --- | --- |
| Master Mix | Enzyme Mix | 100 μL | 2.4 mL |
| 1.2 mL | 1.2 mL |  |  |

- N

- ENote: Store the complete reaction mix on ice or at 4C until use to prevent background signal.

- Preparation of 67% Methanol

- Dilute 100% methanol to 67% using dH2O.

- Standards

Thaw out the 6 tubes containing Histamine Standards. Warm to room temperature (20 – 25°C / 68 – 77°F).

### ELISA Testing Protocol

- Histamine Standards: Add 50 μL of 67% methanol plus 50 μL of each standard (in duplicate) to wells of the 96-well microtiter plate. (Add standards to plate only in the order from low concentration to high concentration). ( Note: If you are using Method II of the Meat/Seafood Extraction Protocol, substitute 50 μL of 1X Enrichment Solution instead of 50 μL 67% methanol).

- Samples: Add 100 μL of each sample in duplicate to wells.

- Add 100 μL of the Reaction Mix to each well.

- Incubate for 10 minutes at room temperature.

- Measure absorbance of each well at 450 nm in a plate reader.

### Histamine Concentration Calculations

A standard curve can be constructed by plotting the mean corrected absorbance obtained from each reference standard against its concentration in ppm.

Calculate the slope and the y-intercept for the line which fits the standard curve data.

The histamine concentration in the well can be described by the equation:

Concentration = (mean absorbance – y-intercept)/slope

Histamine concentration in starting sample = concentration x dilution factor (e.g. 6 for seafood samples).