REAGEN™

EHDV Detection Kit

RNN89006

### Product Description

REAGEN™  EHDV Detection Kit is designed for the detection of Epizootic Hemorrhagic Disease (EHD) virus in cervids. EHDV is an orbivirus that infects most wild and domestic ruminants. EHD is endemic in many countries, particularly affecting white-tailed deer in North America, with a morbidity and mortality rate as high as 90%. Deer, elk, and moose can seroconvert. Edema, hemorrhage, and ulcerations are seen in the mouth and lips in the acute form of the disease. In the chronic form of the disease, affected animals exhibit sloughing of the hoof wall, erosions, and scars in the rumen. Spontaneous abortions and stillbirths have also been reported in some epidemics. Duration of viremia in wild ruminants is 15-40 days post infection, with peak viremia seen from 6 days post infection (titers greater than 106 pfu/ml) in experimental animals

The REAGEN™ EHDV Detection Kit provides a simple, rapid, sensitive and cost-effective enzyme-based immunoassay (ELISA) screening method for detectng EHDV from animal blood sample(serum or plasma) . Bluetongue virus (BTV), a member of the genus Orbivirus, produces similar symptoms, and differential diagnosis is important to prevent widespread outbreak and economic loss. The ELISA methodis specific, without cross-reactivity between the serogroups of BTV and EHDV.

### Procedure OverviewGENERAL INFORMATION

The method described here is a sandwich ELISA. A purified EHDV-specific monoclonal antibody (capture antibody) has been coated in the wells of the microtiter plate. During analysis, the serum or plasma is added. If EHD virus is present in the sample, then it will be captured by the coated antibody. Another monoclonal antibody (detection antibody) that binds to a different EHDV epitope is added to the well. A peroxidase-labeled secondary antibody targets the detection antibody. The resulting color intensity, after addition of substrate, is directly related to the amount of EHDV present in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ EHDV  Detection Kit has the capacity for 96 determinations or testing of 46   samples in duplicate (using 4 wells for negative control and positive control). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. The shelf life is 6 months after receipt when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| EHDV MAb-coated Plate | 1x 96T | 4°C |
| EHDV Positive Control | 0.5 ml | –20°C |
| EHDV Negative Control | 0.5 ml | –20°C |
| 50x Detction MAb | 200 ul | –20°C |
| 50xHRP-Conjugated Antibody Solution | 200 ul | –20°C |
| SecondaryAntibody diluent Solution | 10 ml | 4°C |
| Serum/Plasma Diluent | 15 ml | 4°C |
| 20x Wash Solution | 20 ml | 4°C |
| TMB Substrate | 10 ml | 4°C |
| Stop Solution | 10 ml | 4°C |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Vortex mixer

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette (Optional)

### Warnings and Precautions

REAGEN strongly recommends that you read the following warnings and precautions to ensure your full awareness of the techniques and other details you should pay close attention to when running the assays. Periodically, optimizations and revisions are made to the kit and manual. Therefore, it is important to follow the protocol that comes with the kit. If you need further assistance, you may contact your local distributor.

- Handle the positive and negative controls with care.

- Do not use the kit past the expiration date.

- Do not mix reagents from different kits or lots. Antibodies and plates are lot and kit specific.

- Make sure that the HRP conjugated antibody and diluent are mixed in correct volumes.

- Maintain a laboratory temperature of 20°– 25°C (68°– 77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation.

- Water quality is very important; make sure you are using only distilled or deionized water.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible.

- Add your controls first, and then your samples.

- Always refrigerate plates in sealed bags with desiccant to maintain the plate(s) stability.

- Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25°C / 68 – 77°F) while in the packaging.

### SAMPLE PREPARATION

Serum/Plasma

Dilute serum or plasma samples using the serum/plasma diluent provided in the kit (e.g. 50µL  of serum/plasma is added to 200 µL of diluent to obtain a 1:5 dilution). Perform dilution at room temperature and use immediately. Store any other serum samples at -20°C.

Cell culture Supernatant

Cell culture supernatants from hybridoma or engineered cell lines can be directly used in the assay. Frozen supernatants should be thawed at room temperature before use in the assay.PARATION

INOGANIC PHOSPHORUSDETECTION TEST PCOL

### EHDV DETECTION KIT PROTOCOL

EHDV ANTIBODY ELISA TEST KIT PROTOCOL

Reagent Preparation

IMPORTANT: All frozen reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read the “Warnings and Precautions” section on page 2-3. Solutions should be prepared just prior to performing the ELISA. All reagents should be mixed by inversion prior to use. Do not return used pipette tips or any reagents to the original stock tubes/bottles. Use disposable reservoirs when handling reagents to minimize the risk of contamination.

1. Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Solution with 19 volumes of distilled water.

2. Preparation of 1X Detection Antibody Solution

Based on the needed amount, add 1 volume of 50X Detection Antibody to 49 volumes of the Secondary Antibody Diluent Solution and mix well by inverting. The Secondary Antibody solution is ready to use. The diluted Secondary Antibody is stable for 2 weeks at -20°C.

3. Preparation of 1X HRP-Conjugated Secondary Antibody

Based on the needed amount, add 1 volume of 50X HRP-Conjugated Secondary Antibody solution to 49 volumes of the Secondary Antibody Diluent Solution and mix well by inverting. The Secondary Antibody solution is then ready to use. The diluted Secondary Antibody is stable for 2 weeks at -20°C.

ELISA  Protocol

Amount of reagent required for one reaction is shown in the table below

| Component | Volume per well |
| --- | --- |
| Sample/PositiveControl/Negative Control | 100µL  |
| 1X Detection Antibody Solution | 100µL |
| 1X HRP-Conjugated Antibody | 100µL  |
| 1X Wash Solution | 3.6 mL |
| TMB Substrate | 100µL  |
| Stop Solution | 100 µL |

Note: All samples and controls should be run in duplicate

- Bring the EHDV MAb-coated plate and all reagent components to room temperature (20°– 25°C/ 68°– 77°F) for at least an hour.

- Aliquot EHDV Positive Control (100µL/well) to two wells of the antigen coated plate.

- Aliquot EHDV Negative Control (100 µL/well) to two wells of the antigen coated plate.

Use extreme care while aliquoting the above controls. Change to a clean pipette tip each time liquid is withdrawn.

- Add 100µL of previously diluted (1:5) serum/plasma sample (or neat cell culture supernatant) per well in duplicate.

- Cover plate with foil and incubate for 60 minutes at 37°C. Avoid direct sunlight and air vents during incubation.

- Discard the solution in the wells.

- Wash the plate by adding 300µL of 1X Wash solution to each well of the plate. Discard the wash solution and tap dry on paper towels. Repeat three additional times for a total of four washes. Perform the next step immediately. Do not allow the plate to dry.

- Add 100µL of the 1X Detection Antibody to each well of the plate. Cover plate with foil and incubate for 60 minutes at room temperature (20 – 25°C / 68 – 77°F).

- Wash plate with 1X Wash solution as described in Step 7.

- Add 100µL of 1X HRP-Conjugated Secondary Antibody to each well of the plate. Cover plate with foil and incubate for 60 minutes at 37°C.

- Wash plate with 1X Wash solution as described in Step 7.

- Add 100µL of TMB Substrate to each well of the plate. Cover plate with foil and incubate for 5 minutes at room temperature (20 – 25°C / 68 – 77°F).

- Add 100µL of Stop Solution to each well. Read the plate as soon as possible on a plate reader at 450nm wavelength.

RESULTS:

Interpretation of Controls

| Controls | OD(450) | EHDV titer |
| --- | --- | --- |
| Positive Control: | ≤0.2 | EHDV negative |
| Negative Control: | Three times OD450 of negative control | ≥105 pfu/mL |

pfu/mL: plaque-forming units per mL as assayed by infection on susceptible cell lines.

Controls must be within the limits above for proper interpretation of samples.

Interpretation of Samples

Sample O.D450 ≥ 3x Negative Control OD450

Sample is interpreted as positive for EHDV with a titer ≥ 105 pfu/mL.

Sample O.D450: ≥ 1X Negative Control O.D450 ≤ 3X

Sample needs to be retested and/or confirmed for presence of EHDV by other methods.

Sample O.D450: ≤ Negative Control O.D450

Sample is interpreted as negative for EHDV

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com    reagenllc@gmail.com

Web.:www.reagen.us