REAGEN™

Boldenone ELISA Test Kit

RNH 94008

Product Description

REAGEN Boldenone ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Boldenone in fish,shrimp, meat and urine. The unique features of the kit are:

- High recovery (>80%),rapid(10-40 minutes),and cost-effective extraction methods.

- High sensitivity(0.05 ppb)and low detection limit(0.025ppb) for meat sample

- A quick ELISA assay(less than 2 hours regardless of number of samples).

- High reproducibility.

Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target toxin and enzyme-conjugate. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The enzyme-conjugate, will bind with the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

Kit Contents, Storage and Shelf Life

REAGEN Boldenone ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8°C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Boldenone Coated Plate | 1 x 96-well Plate (8 wells x 12 strips) | 2-8°C |
| Boldenone Standards: Negative control (white cap tube) 0.05 ng/mL (yellow cap tube) 0.15 ng/mL (orange cap tube) 0.5 ng/mL (pink cap tube) 1.5 ng/mL (purple cap tube) 4.5 ng/mL (blue cap tube) 10 µg/mL (spiking, optional, red cap tube) | 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL | 2-8°C * |
| Boldenone Antibody | 8 mL | 2-8°C * |
| 100X HRP Conjugate | 250 µL | 2-8°C * |
| HRP Conjugate Diluent | 20 mL | 2-8°C |
| 20X Wash Solution | 28 mL | 2-8°C |
| Stop Buffer | 20 mL | 2-8°C |
| TMB Substrate | 12 mL | 2-8°C |
| 10 X PBS | 25 mL | 2-8°C |
| Sample Balance Buffer A | 10 mL | 2-8°C |
| Sample Balance Buffer B | 62g | 2-8°C |
| Tissue Clean Up mix | 25g | 2-8°C |

*	If you are not planning to use the kit for over 1 month, store  Boldenone     Standards, Boldenone Antibody and  100X HRP Conjugate at -20°C or in a freezer.

Sensitivity (Detection Limit)

| Sample Type | Detection Limit (g/g or ppb) |
| --- | --- |
| Meat/Fish/shrimp | 0.025 |
| Urine | 0.25 |

Specificity (Cross-Reactivity)

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Boldenone | 100 |
| Nandrolone | 5 |
| Testosterone | 12 |
| Mesterolone | <0.5 |
| Androsterone | 0.1 |
| Trenbolone | <0.1 |

Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Rotary evaporator or nitrogen gas

- Vortex mixer (e.g. Genie Vortex mixer from VWR)

- 10, 20, 100 and 1000 µL pipettes

- Multi-channel pipette: 50-300 µL (Optional)

Warnings and Precautions

- The standards contain Boldenone. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within  their  expiration  dates. ANTIBODIES  AND PLATES  ARE  KIT-AND LOT-SPECIFIC.  Make sure that the HRP Conjugate and Diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays  in  direct sunlight,  as this may cause excessive  heat and evaporation. Cold  bench tops should be avoided by placing several layers of paper towel or some other insulate material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very

important.

- When pipetting samples or reagents  into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards  first and then your samples.

- Add standards  to plate only in the order from low concentration to high concentration, as this will minimize the risk of compromising the standard curve.

- Always  refrigerate  plates  in  sealed  bags  with  a  desiccant  to  maintain  stability.  Prevent condensation from forming on plates by allowing them equilibrate to room temperature  (20 – 25°C /68 – 77°F) while in the packaging.

# SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C  if  they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25°C / 68 – 77°F) or in a refrigerator before use.

1. Preparation of 1X PBS

Mix 1 volume of the 10 X PBS with 9 volumes of distilled water.

2. Preparation of 1X Boldenone Sample balance Buffer

Add 10 mL of the Sample Balance Buffer I and 62 g of Sample Balance Buffer II (Powder)  to a 250-mL plastic bottle. Add 190 mL of distilled water to the bottle, mix well to obtain a  homogeneous solution.

Fish/Shrimp/Meat (Beef/Chicken/Pork)

1.  Weigh 2 g homogenized tissue sample, add 6 mL of acetonitrile and 2.5 mL of 1X Sample Balance Buffer

2.  Vortex for 3 minutes at maximum speed or shake for 20 minutes at a  multi-tube vortexer or shaker.

3.  Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25°C       / 68 – 77°F), transfer 3 mL of the supernatant to a new tube.

4.  Add 300 mg Tissue Clean Up Mix and vortex for 30 seconds, leave at room temp for 5 minutes.

5.  Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25°C       / 68 – 77°F),

6.  Transfer 2.4 mL of supernatant to a new tube. Use a rotary evaporator      to dry the sample in a 50-60°C water bath under reduced pressure.   Alternatively, the sample can be dried by  blowing nitrogen gas in a 50-60°C  water bath.

7.  Add 300 µL of 1X PBS, vortex vigorously for 30 seconds at maximum speed.

8.  Use 50 µL of the sample for the assay.

Note: Dilution factor: 0.5.

Urine

1.  Centrifuge 0.5 mL of the urine sample at 4,000 x g for 5 minutes.

2.  Take out 0.1 mL of supernatant, add 0.4 mL of 1X PBS, mix well.

3.  Use 50 µL per well for the assay.

Note: Dilution factor: 5.

# BOLDENONE ELISA TEST KIT PROTOCOL

Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at20 – 25°C / 68 – 77°F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.    All reagents should be mixed by gently inverting or swirling prior to use. Prepare only the volumes that are needed for the number of wells being run. Do  not return the  reagents  to the  original stock  tubes/bottles. It  is recommended  that disposable reservoirs be used when handling reagents to minimize the risk of contamination.

1.	Preparation of 1X HRP Conjugate

Mix 1 volume of 100X HRP Conjugate with 99 volumes of HRP ConjugateDiluent .

- Preparation of 1X Wash Solution

Mix 1 volume of the 20 X Wash Solution  with 19 volumes of distilled water.

ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Boldenone Antibody | 100 µL | 2.4 mL |
| 1X HRP Conjugate | 150 µL | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 µL | 2.4 mL |
| TMB Substrate | 100 µL | 2.4 mL |

1.   A Add 50 µL of each Boldenone Standard in duplicate into different wells ( Add standards to plate  only in the order from low concentration to high concentration).

2.  Add 50 µL of each sample in duplicate into different sample wells.

3.  Add 100 µL of Antibody #1 and mix well by gently rocking the plate manually for 1 minute.

4.  Incubate the plate for 30 minutes at room temperature (20 – 25°C / 68 – 77°F).

5.  Wash the plate 3 times with 250 µL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (  Perform the next step immediately after plat

washings. Do not allow the plate to air dry between working steps)

6.  Add 150 µL of 1X Antibody #2 solution. Incubate the plate for 30 minutes at room temperature (20 – 25°C / 68 – 77°F) (    Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

7.  Wash the plate 3 times with 250 µL 1X Wash Solution. After the last wash, invert the plate and

gently tap the plate dry on paper towels (  Perform the next step immediately after plate  washings. Do not allow the plate to air dry between working steps).

8.  Add 100 µL of TMB substrate. Start timing the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating at room temperature ( Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and  should  be  discarded.  Covering  the  microtiter  plate  while  incubating recommended).

9.  After incubating for 15 minutes at room temperature (20 – 25°C / 68 – 77°F), add 100 µL of

Stop Buffer to stop the enzyme reaction.

10. Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength (  Before reading, use a lint-free wipe on the bottom of the plate to ensure

no moisture or fingerprints interfere with the readings).

Boldenone Concentration Calculations

A standard  curve can be constructed by plotting the mean relative absorbance  (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of  the tested  drug in ng/mL from the  standard  curve.

The following figure is a typical Boldenone  standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0251         FAX.:856-727-0250

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us